# OWH67E Series

# Programmable DC Power Supply Series

# Features

● 1/2 2U chassis with rear outputs for rack integration.

● Wide-range AC input and 1,000 W constant-power output, with 60 A/120 A models ideal for low-voltage, high-current testing.

● 10 mV/1 mA setting and readback resolution.

● 100-step LIST programming with PC control.

● CV/CC priority switching for different load requirements.

● Built-in discharge circuit removes residual voltage after shutdown for safer operation.

● Built-in SAS model supports EN 50530 and Sandia PV curves for MPPT efficiency testing.

● Digital control with OVP, OCP, OPP, OTP, and other protections.

● Standard RS232/RS485 interfaces support the Modbus protocol and SCPI commands for remote control and automated testing.

● 3.9-inch TFT color LCD with output curves and 1-minute data logging.

Specifications

Specifications are guaranteed after a 30-minute warm-up at the specified temperature.

| Model AC Input (Voltage/Frequency) |  | OWH67E010-150 OWH67E010-60 |  |
| --- | --- | --- | --- |
|  |  | 100-240Vac；50-60Hz |  |
| Rated Output (0 °C to 40 °C) | Voltage | 0-150V | 0-60V |
|  | OVP | 155V | 65V |
|  | Current | 0-60A | 0-120A |
| Maximum Output Power |  | 1000W |  |
| Load Regulation* | Voltage | ≤30mV | ≤30mV |
|  | Current | ≤30mA | ≤30mA |
| Line Regulation | Voltage | ≤20mV | ≤20mV |
|  | Current | ≤20mA | ≤20mA |
| Settings Resolution |  | 10mV/1mA |  |
| Readback Resolution |  | 10mV/1mA |  |
| Settings Accuracy (% of output + offset) | Voltage | ≤0.1%+0.1%FS | ≤0.1%+0.1%FS |
|  | Current | ≤0.1%+0.1%FS | ≤0.1%+0.1%FS |
| Readback Accuracy (% of output + offset) | Voltage | ≤0.1%+0.1%FS | ≤0.1%+0.1%FS |
|  | Current | ≤0.1%+0.1%FS | ≤0.1%+0.1%FS |
| Ripple & Noise (20 Hz to 20 MHz) | Voltage | ≤ 150mVp-p | ≤ 150mVp-p |
|  | Current(rms) | ≤60mA | ≤ 120mA |
| Output Temperature Coefficient (0℃-40℃) |  | 100 ppm/°C (Voltage); 200 ppm/°C (Current) |  |
| Readback Temperature Coefficient |  | 100 ppm/°C (Voltage); 200 ppm/°C (Current) |  |
| Recovery Time (10% to 90% Load) |  | ≤ 5.0ms |  |
| Display |  | 3.9-inch TFT Color LCD |  |
| Interface |  | RS232, RS485, and CAN Parallel Operation |  |
| Dimensions (W × H × D, mm) |  | 214*88*374 |  |
| Weight (kg) |  | Approx. 6.3 kg |  |

*: Voltage load regulation is measured at the remote sense terminals.                      Specifications subject to change without prior notice.

Ordering Information

| Ordering Information |  |  | Order No. |
| --- | --- | --- | --- |
| Main Unit |  |  |  |
| 0-150V ，0-60A， | 1000W， | 1CH | OWH67E010-150 |
| 0-60V ，0-120A， | 1000W， | 1CH | OWH67E010-60 |
| Standard Accessories |  |  |  |
| Quick Guide |  |  | ———— |
| Power Cord (Local Standard) |  |  | ———— |
| RS485/CAN Communication Terminal Block |  |  | 7Term-3 |
| Remote Sense Terminal Block |  |  | 7Term-4 |
| Optional Accessories |  |  |  |
| 40A Test Leads |  |  | 7T2T40 |
| 80A Test Leads |  |  | 7T2T80 |
| 120A Test Leads |  |  | 7T2T120 |
| L-Shaped Mounting Bracket |  |  | 7L-B |
| Warranty |  |  |  |
| 1-year main-unit warranty; accessories excluded. |  |  |  |

Fujian Lilliput Optoelectronics Technology Co., Ltd.

No. 19, Heming Road, Lantian Industrial Zone Zhangzhou 363005 P.R. China Tel : +86.596.213.0430         E-mail: info@owon.com.cn         www.owon.com