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Programmable DC Power Supply Series
Features
● 1/2 2U chassis with rear outputs for rack integration.
● Wide-range AC input and 1,000 W constant-power output, with 60 A/120 A models ideal for low-voltage, high-current testing.
● 10 mV/1 mA setting and readback resolution.
● 100-step LIST programming with PC control.
● CV/CC priority switching for different load requirements.
● Built-in discharge circuit removes residual voltage after shutdown for safer operation.
● Built-in SAS model supports EN 50530 and Sandia PV curves for MPPT efficiency testing.
[image: ]● Digital control with OVP, OCP, OPP, OTP, and other protections.
● Standard RS232/RS485 interfaces support the Modbus protocol and SCPI commands for remote control and automated testing.
● 3.9-inch TFT color LCD with output curves and 1-minute data logging.
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[image: ] Specifications
Specifications are guaranteed after a 30-minute warm-up at the specified temperature.
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AC Input (Voltage/Frequency)
	 (
OWH67E010-150
)[image: ]  OWH67E010-60

	
	100-240Vac；50-60Hz

	Rated Output (0 °C to 40 °C)
	Voltage
	0-150V
	0-60V

	
	OVP
	155V
	65V

	
	Current
	0-60A
	0-120A

	Maximum Output Power
	1000W

	Load Regulation*
	Voltage
	≤30mV
	≤30mV

	
	Current
	≤30mA
	≤30mA

	Line Regulation
	Voltage
	≤20mV
	≤20mV

	
	Current
	≤20mA
	≤20mA

	Settings Resolution
	10mV/1mA

	Readback Resolution
	10mV/1mA

	Settings Accuracy (% of output + offset)
	Voltage
	≤0.1%+0.1%FS
	≤0.1%+0.1%FS

	
	Current
	≤0.1%+0.1%FS
	≤0.1%+0.1%FS

	Readback Accuracy
(% of output + offset)
	Voltage
	≤0.1%+0.1%FS
	≤0.1%+0.1%FS

	
	Current
	≤0.1%+0.1%FS
	≤0.1%+0.1%FS

	Ripple & Noise
(20 Hz to 20 MHz)
	Voltage
	≤ 150mVp-p
	≤ 150mVp-p

	
	Current(rms)
	≤60mA
	≤ 120mA

	Output Temperature Coefficient (0℃-40℃)
	100 ppm/°C (Voltage); 200 ppm/°C (Current)

	Readback Temperature Coefficient
	100 ppm/°C (Voltage); 200 ppm/°C (Current)

	Recovery Time (10% to 90% Load)
	≤ 5.0ms

	Display
	3.9-inch TFT Color LCD

	Interface
	RS232, RS485, and CAN Parallel Operation

	Dimensions (W × H × D, mm)
	214*88*374

	Weight (kg)
	Approx. 6.3 kg


*: Voltage load regulation is measured at the remote sense terminals.                      Specifications subject to change without prior notice.
[image: ] Ordering Information

	Ordering Information
	Order No.

	Main Unit

	0-150V ，0-60A，
	1000W，
	1CH
	OWH67E010-150

	0-60V ，0-120A，
	1000W，
	1CH
	OWH67E010-60

	Standard Accessories

	Quick Guide
	————

	Power Cord (Local Standard)
	————

	RS485/CAN Communication Terminal Block
	7Term-3

	Remote Sense Terminal Block
	7Term-4

	Optional Accessories

	40A Test Leads
	7T2T40

	80A Test Leads
	7T2T80

	120A Test Leads
	7T2T120

	L-Shaped Mounting Bracket
	7L-B

	Warranty

	1-year main-unit warranty; accessories excluded.
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