OEL15&30 Series Technical Specifications

The instrument must be operated continuously for more than 30 minutes at the specified operating temperature to achieve the following specifications:

| Model |  | OEL 1530 | OEL 1520 | OEL 1515 | OEL 3030 | OEL 3020 | OEL 3015 | OEL 6030 | OEL 6020 | OEL 6015 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Input | Voltage | 0-150V |  |  | 0-300V |  |  | 0-600V |  |  |
|  | Current | 0-30A | 0-20A | 0-15A | 0-30A | 0-20A | 0-15A | 0-30A | 0-20A | 0-15A |
|  | Power | 300W |  | 150W | 300W |  | 150W | 300W |  | 150W |
| Minimum Operation voltage |  | 1.1V@30A | 0.9V@20A | 0.8V@15A | 1.1V@30A | 0.9V@20A | 0.8V@15A | 1.1V@30A | 0.9V@20A | 0.8V@15A |
| Constant voltage mode (CV) |  |  |  |  |  |  |  |  |  |  |
| Range |  | 0-150V |  |  | 0-300V |  |  | 0-600V |  |  |
| Resolution |  | 1mV |  |  |  |  |  |  |  |  |
| Accuracy |  | 0.05%+0.05%FS |  |  | 0.05%+0.05%FS |  |  | 0.05%+0.05%FS |  |  |
| Constant current mode (CC) |  |  |  |  |  |  |  |  |  |  |
| Range |  | 0-30A | 0-20A | 0-15A | 0-30A | 0-20A | 0-15A | 0-30A | 0-20A | 0-15A |
| Resolution |  | 1mA |  |  | 1mA |  |  | 1mA |  |  |
| Accuracy |  | 0.05%+0.05%IFS |  |  | 0.05%+0.05%IFS |  |  | 0.05%+0.05%IFS |  |  |
| Constant resistance mode (CR) |  |  |  |  |  |  |  |  |  |  |
| Range |  | 0.05-7.5k Ω |  |  | 0.05-7.5k Ω |  |  | 0.05-7.5k Ω |  |  |
| Accuracy*1 |  | 0.1%+0.01%R |  |  | 0.1%+0.01%R |  |  | 0.1%+0.01%R |  |  |
| Constant power mode (CP) |  |  |  |  |  |  |  |  |  |  |
| Range |  | 300W |  | 150W | 300W |  | 150W | 300W |  | 150W |
| Resolution |  | 10mW |  |  | 10mW |  |  | 10mW |  |  |
| Accuracy*1 |  | 0.1%+0.1%FS |  |  | 0.1%+0.1%FS |  |  | 0.1%+0.1%FS |  |  |
| Dynamic mode (CCD) |  |  |  |  |  |  |  |  |  |  |
| T1&T2 |  | 0.1ms-50s |  |  | 0.1ms-50s |  |  | 0.1ms-50s |  |  |
| Rise/Fall slope |  | 0.01A/ms-2000A/ms |  |  | 0.01A/ms-2000A/ms |  |  | 0.01A/ms-2000A/ms |  |  |
| Min rise time |  | 20us |  |  | 20us |  |  | 20us |  |  |
| Voltage readback value |  |  |  |  |  |  |  |  |  |  |
| Range |  | 0-15V/ 0-150V |  |  | 0-30V/ 0-300V |  |  | 0-60V/ 0-600V |  |  |
| Resolution |  | 1mV |  |  | 1mV |  |  | 1mV |  |  |
| Accuracy |  | 0.05%+0.05%FS |  |  | 0.05%+0.05%FS |  |  | 0.05%+0.05%FS |  |  |
| Current readback value |  |  |  |  |  |  |  |  |  |  |
| Range |  | 0-30A | 0-20A | 0-15A | 0-30A | 0-20A | 0-15A | 0-30A | 0-20A | 0-15A |
| Resolution |  | 0.1mA |  |  | 0.1mA |  |  | 0.1mA |  |  |
| Accuracy |  | 0.05%+0.05%FS |  |  | 0.05%+0.05%FS |  |  | 0.05%+0.05%FS |  |  |
| Power readback value |  |  |  |  |  |  |  |  |  |  |
| Range |  | 300W |  | 150W | 300W |  | 150W | 300W |  | 150W |
| Resolution |  | 10mW |  |  | 10mW |  |  | 10mW |  |  |
| Accuracy |  | 0.1%+0.1%FS |  |  | 0.1%+0.1%FS |  |  | 0.1%+0.1%FS |  |  |
| Protection function |  |  |  |  |  |  |  |  |  |  |
| OPP |  | >OPP value 10ms protection |  |  | >OPP value 10ms protection |  |  | >OPP value 10ms protection |  |  |
| OVP |  | >OVP value 0.1ms protection |  |  | >OVP value 0.1ms protection |  |  | >OVP value 0.1ms protection |  |  |
| OCP |  | >OCP value 0.1ms protection |  |  | >OCP value 0.1ms protection |  |  | >OCP value 0.1ms protection |  |  |
| OTP |  | 85℃ |  |  | 85℃ |  |  | 85℃ |  |  |
| Short circuit |  |  |  |  |  |  |  |  |  |  |
| Current (CC) |  | >30A | >20A | >15A | >30A | >20A | >15A | >30A | >20A | >15A |
| Voltage (CV) |  | 0 |  |  | 0 |  |  | 0 |  |  |
| Resistance (CR) |  | 50mΩ |  |  | 50mΩ |  |  | 50mΩ |  |  |
| Test Function |  |  |  |  |  |  |  |  |  |  |
| Test Function |  | Battery test, OCP,OPP, Optional, only available for “T”version Ripple test function, Optional, only available for “V”version |  |  |  |  |  |  |  |  |
| AC parameter |  |  |  |  |  |  |  |  |  |  |
| Voltage range |  | 100V-240Vac |  |  |  |  |  |  |  |  |
| Frequency range |  | 50/60Hz |  |  |  |  |  |  |  |  |
| Interface |  |  |  |  |  |  |  |  |  |  |
| RS-232 /RS458 |  | Support RS232, RS485, USB, Supported SCPI |  |  |  |  |  |  |  |  |
| General |  |  |  |  |  |  |  |  |  |  |
| Heat dissipation |  | Internal intelligent fan forced air cooling |  |  |  |  |  |  |  |  |
| Withstand voltage (Input to Ground) |  | 1.5KV/AC/5mA |  |  |  |  |  |  |  |  |
| Volume (L*W*H) |  | 282.9mm*84.3mm*151.15mm |  |  |  |  |  |  |  |  |
| Net weight |  | Approx. 2.5kg |  |  | Approx. 2.5kg |  |  | Approx. 2.5kg |  |  |

* 1 Accuracy: The voltage and current input values should not be less than 10%FS.

If there are any changes to parameters, no separate notice will be given. Please refer to the website update for the latest information.

Interval Period of Adjustment: One year is recommended for the calibration interval period.

V1.0.1