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The instrument must be operated continuously for more than 30 minutes at the specified operating temperature to achieve the following specifications:
	Model
	OEL
1530
	OEL
1520
	OEL
1515
	OEL
3030
	OEL
3020
	OEL
3015
	OEL
6030
	OEL
6020
	OEL
6015

	Input
	Voltage
	0-150V
	0-300V
	0-600V

	
	Current
	0-30A
	0-20A
	0-15A
	0-30A
	0-20A
	0-15A
	0-30A
	0-20A
	0-15A

	
	Power
	300W
	150W
	300W
	150W
	300W
	150W

	Minimum Operation voltage
	1.1V@30A
	0.9V@20A
	0.8V@15A
	1.1V@30A
	0.9V@20A
	0.8V@15A
	1.1V@30A
	0.9V@20A
	0.8V@15A

	Constant voltage mode (CV)

	Range
	0-150V
	0-300V
	0-600V

	Resolution
	1mV

	Accuracy
	0.05%+0.05%FS
	0.05%+0.05%FS
	0.05%+0.05%FS

	Constant current mode (CC)

	Range
	0-30A
	0-20A
	0-15A
	0-30A
	0-20A
	0-15A
	0-30A
	0-20A
	0-15A

	Resolution
	1mA
	1mA
	1mA

	Accuracy
	0.05%+0.05%IFS
	0.05%+0.05%IFS
	0.05%+0.05%IFS

	Constant resistance mode (CR)

	Range
	0.05-7.5k Ω
	0.05-7.5k Ω
	0.05-7.5k Ω

	Accuracy*1
	0.1%+0.01%R
	0.1%+0.01%R
	0.1%+0.01%R

	Constant power mode (CP)

	Range
	300W
	150W
	300W
	150W
	300W
	150W

	Resolution
	10mW
	10mW
	10mW

	Accuracy*1
	0.1%+0.1%FS
	0.1%+0.1%FS
	0.1%+0.1%FS

	Dynamic mode (CCD)

	T1&T2
	0.1ms-50s
	0.1ms-50s
	0.1ms-50s

	Rise/Fall slope
	0.01A/ms-2000A/ms
	0.01A/ms-2000A/ms
	0.01A/ms-2000A/ms

	Min rise time
	20us
	20us
	20us

	Voltage readback value

	Range
	0-15V/
0-150V
	0-30V/
0-300V
	0-60V/
0-600V

	Resolution
	1mV
	1mV
	1mV

	Accuracy
	0.05%+0.05%FS
	0.05%+0.05%FS
	0.05%+0.05%FS

	Current readback value

	Range
	0-30A
	0-20A
	0-15A
	0-30A
	0-20A
	0-15A
	0-30A
	0-20A
	0-15A

	Resolution
	0.1mA
	0.1mA
	0.1mA

	Accuracy
	0.05%+0.05%FS
	0.05%+0.05%FS
	0.05%+0.05%FS

	Power readback value

	Range
	300W
	150W
	300W
	150W
	300W
	150W

	Resolution
	10mW
	10mW
	10mW

	Accuracy
	0.1%+0.1%FS
	0.1%+0.1%FS
	0.1%+0.1%FS

	Protection function

	OPP
	>OPP value 10ms protection
	>OPP value 10ms protection
	>OPP value 10ms protection

	OVP 
	>OVP value 0.1ms protection
	>OVP value 0.1ms protection
	>OVP value 0.1ms protection

	OCP
	>OCP value 0.1ms protection
	>OCP value 0.1ms protection
	>OCP value 0.1ms protection

	OTP
	85℃
	85℃
	85℃

	Short circuit

	Current (CC)
	>30A
	>20A
	>15A
	>30A
	>20A
	>15A
	>30A
	>20A
	>15A

	Voltage (CV)
	0
	0
	0

	Resistance (CR)
	50mΩ
	50mΩ
	50mΩ

	Test Function

	Test Function
	Battery test, OCP,OPP, Optional, only available for “T”version
Ripple test function, Optional, only available for “V”version

	AC parameter

	Voltage range
	100V-240Vac

	Frequency range
	50/60Hz

	Interface

	RS-232 /RS458
	Support RS232, RS485, USB, Supported SCPI

	General

	Heat dissipation
	Internal intelligent fan forced air cooling

	Withstand voltage (Input to Ground)
	1.5KV/AC/5mA

	Volume (L*W*H)
	282.9mm*84.3mm*151.15mm

	Net weight 
	Approx. 2.5kg
	Approx. 2.5kg
	Approx. 2.5kg


* 1 Accuracy: The voltage and current input values should not be less than 10%FS.
If there are any changes to parameters, no separate notice will be given. Please refer to the website update for the latest information.
Interval Period of Adjustment: One year is recommended for the calibration interval period.
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