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OEL15&30&60 Series Programmable DC Electronic Load 

User Manual
For product support, visit:www.owon.com.hk/download
※:The illustrations, interface, icons and characters in the user manual may be slightly

different from the actual product. Please refer to the actual product.
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Copyright © LILLIPUT Company. All rights reserved. 
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 is the registered trademark of the LILLIPUT Company. 

Fujian LILLIPUT Optoelectronics Technology Co., Ltd. 

No. 19, Heming Road 

Lantian Industrial Zone, Zhangzhou 363005 P.R. China
Tel: +86-596-2130430            Fax: +86-596-2109272 

Web: www.owon.com.cn          E-mail: info@owon.com.cn
General Warranty

We warrant that the product will be free from defects in materials and workmanship for a period of 2 years from the date of purchase of the product by the original purchaser from our company. The warranty period for accessories are 12 months. This warranty only applies to the original purchaser and is not transferable to a third party.
If the product proves defective during the warranty period, we will either repair the defective product without charge for parts and labour, or will provide a replacement in exchange for the defective product. Parts, modules and replacement products used by our company for warranty work may be new or reconditioned like new. All replaced parts, modules and products become the property of our company.
To obtain service under this warranty, the customer must notify our company of the defect before the expiration of the warranty period. Customer shall be responsible for packaging and shipping the defective product to our designated service centre, a copy of the customer's proof of purchase is also required.

This warranty shall not apply to any defect, failure or damage caused by improper use or improper or inadequate maintenance and care. We shall not be obligated to furnish service under this warranty a) to repair damage resulting from attempts by personnel other than our company representatives to install, repair or service the product; b) to repair damage resulting from improper use or connection to incompatible equipment; c) to repair any damage or malfunction caused by the use of not our supplies; or d) to service a product that has been modified or integrated with other products when the effect of such modification or integration increases the time or difficulty of servicing the product. 
Please contact the nearest sales and service offices for services.

Excepting the after-sales services provided in this summary or the applicable warranty statements, we will not offer any guarantee for maintenance definitely declared or hinted, including but not limited to the implied guarantee for marketability and special-purpose acceptability. We should not take any responsibilities for any indirect, special or consequent damages.
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General Safety Requirements
Before use, please read the following safety precautions to avoid any possible bodily injury and to prevent this product or any other connected products from damage. In order to avoid any contingent danger, ensure this product is only used within the range specified. 

Only the qualified technicians can implement the maintenance. 

To avoid Fire or Personal Injury:

Use Proper Power Cord. Use only the power cord supplied with the product and certified to use in your country.

Connect the probe correctly. The grounding end of the probe corresponds to the grounding phase. Please don't connect the grounding end to the positive phase.
Check all Terminal Ratings. To avoid fire or shock hazard, check all ratings and markers of this product. Refer to the user's manual for more information about ratings before connecting to the instrument.

Do not operate without covers. Do not operate the instrument with covers or panels removed.

Use Proper Fuse. Use only the specified type and rating fuse for this instrument.

Avoid exposed circuit. Do not touch exposed junctions and components when the instrument is powered.

Do not operate if in any doubt. If you suspect damage occurs to the instrument, have it inspected by qualified service personnel before further operations.

Use your instrument in a well-ventilated area. Make sure the instrument installed with proper ventilation, refer to the user manual for more details.
Do not operate in wet conditions.To avoid the risk of internal circuit short circuits or electric shock, do not operate the instrument in a humid environment.
Do not operate in an explosive atmosphere.To avoid damage to the instrument or personal injury, do not operate the instrument in flammable or explosive environments.

Keep product surfaces clean and dry.To prevent dust or moisture in the air from affecting the instrument's performance, please keep the product's surface clean and dry.

Safety Terms And Symbols 

Safety Terms
Terms in this manual. The following terms may appear in this manual: 
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  Warning: Warning indicates the conditions or practices that could result in injury or loss of life. 
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  Caution: Caution indicates the conditions or practices that could result in damage to this product or other property. 
Terms on the product. The following terms may appear on this product: 
Danger: It indicates an injury or hazard may immediately happen. 

Warning: It indicates an injury or hazard may be accessible potentially. 

Caution: It indicates a potential damage to the instrument or other property might occur.
Safety Symbols

Symbols on the product. The following symbol may appear on the product:
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	Hazardous Voltage
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	Caution, risk of danger (refer to this manual for specific Warning or Caution information)
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	Protective Earth Terminal
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	Chassis Ground
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	Test Ground
	
	


Quick Start

This chapter will introduce the power-on inspection steps for the load, ensuring that the load can be properly turned on and used in its initialized state. It will also cover the front panel, rear panel, button functions, and display functions of the load, ensuring that before operating the load, users can quickly understand the load's appearance, structure, and button functionality, which will help you make better use of this series of loads.

Product Introduction

The single-channel programmable electronic load, it is primarily designed for applications such as fast chargers, battery discharge, and power supply testing, and other product performance validation tests. The series provides optimal solution for design R&D and production line testing.

Model Selection Table:

	Model
	Voltage
	Current
	Power
	Communication interface

	OEL1515
	150V
	15A
	150W
	USB, RS232, RS485

	OEL1520
	150V
	20A
	300W
	USB, RS232, RS485

	OEL1530
	150V
	30A
	300W
	USB, RS232, RS485

	OEL3015
	300V
	15A
	150W
	USB, RS232, RS485

	OEL3020
	300V
	20A
	300W
	USB, RS232, RS485

	OEL3030
	300V
	30A
	300W
	USB, RS232, RS485

	OEL6015
	600V
	15A
	150W
	USB, RS232, RS485

	OEL6020
	600V
	20A
	300W
	USB, RS232, RS485

	OEL6030
	600V
	30A
	300W
	USB, RS232, RS485


Panel and User Interface

Front Panel
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Figure 3‑1 Front Panel

	1
	Display
	Displays the user interface.

	2
	F1-F3
	Submenu option setting keys.

	3
	SHIFT
	Multi-function key.

	4
	CV/CP
	Constant voltage mode shortcut key: The user enters the constant voltage setting mode;

SHIFT + this key: Constant power mode shortcut key, allowing the user to enter the constant power setting interface.

	5
	Knob
	The user can rotate the knob to input programming information or select options.

	6
	SHORT
	Short circuit mode.

	7
	On/Off
	Controls the output state: ON or OFF.

SHIFT + this key: Multi-function, LOCK function.

	8
	ENTER
	Confirmation key.

	9
	USB
	Type C USB data interface.

	10
	Channel Output Terminal
	Output connection of the channel.

	11
	PE
	Earth terminal on the machine chassis.

	12
	ESC
	Cancel editing or go back to the previous menu.

	13
	Vsense-
	Remote sensing, negative polarity terminal.

	14
	TRIG/TRAN
	Trigger function key;

SHIFT + this key: Multi-function, allows the user to enter dynamic mode.

	15
	Vsense+
	Remote sensing, positive polarity terminal.

	16
	CC/CR
	Constant current mode shortcut key: The user enters the constant current setting mode;

SHIFT + this key: Constant resistance mode shortcut key, allowing the user to enter the constant resistance setting interface.


Rear Panel
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	Figure 3‑2 Rear Panel

	1
	Air Vent
	Do not block the ventilation outlet, as it will prevent the machine from dissipating heat properly, which could lead to overheating inside the machine.

	2
	Power Button
	Turn On/Off the instrument.

	3
	Fuse
	Power fuse.

	4
	AC Power Input Jack
	The power cable inputs AC power from this connection terminal to the input terminal.

	5
	DB9 Terminal
	Includes: RS232, RS485, and Trig In/Out; communication control can be done via RS232 and RS485. (Note: Different models are equipped with different communication interfaces as options.)


User Interface
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Figure 3-3 User Interface

Status Icon

	Icon
	Description
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	The panel keys are locked
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	Enable beeper
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	A failure alarm


General Inspection
When you receive a new DC electronic load, it is recommended that you follow the steps below to inspect the instrument.

1. Check whether there is any damage caused by transportation.
If it is found that the packaging carton or the foamed plastic protection cushion has suffered serious damage, do not throw it away first till the complete device and its accessories succeed in the electrical and mechanical property tests.
2. Check the Accessories.
The supplied accessories have been already described in Appendix A: Accessories of this manual. You can check whether there is any loss of accessories with reference to this description. If it is found that there is any accessory lost or damaged, please get in touch with our distributor responsible for this service or our local offices.
3．Check the Complete Instrument.
If it is found that there is damage on the first appearance of the instrument, or the instrument cannot work normally, or fails in the performance test, please get in touch with our distributor responsible for this business or our local offices. If there is damage to the instrument caused by the transportation, please keep the package. With the transportation department or our distributor responsible for this business informed about it, a repairing or replacement of the instrument will be arranged by us.
Power-on Inspection
Connect the instrument to an AC power source using the power cord supplied with the accessory.
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	Warning:
To prevent electric shock, make sure the instrument is properly grounded.


Press the power button on the front panel and the startup screen will be displayed on the screen.
Functional features and Panel operation
This chapter will provide a detailed description of the features and functions of the DC electronic load, as well as the panel menu operations. It will be divided into the following sections:

Turn On/Off channel output

Local/Remote mode operation switch

Constant mode function

Input control function

Keyboard lock function

Short circuit simulation function

System menu function

Multiple device parallel function

Fault information

Introduction to communication interfaces

Test function

Turn On/Off the Channel Output

Press the On/Off key to turn On/Off the channel. When the  On/Off key is bright, the output is turned ON; when the On/Off key light is off, the output is turned off. When the device is turned on, the panel defaults to the main interface to display the current state.

Local/Remote mode Operation switch

The instrument provides both local and remote operation mode, and the two modes can be switched by communication commands. In power initialization mode, the machine is operated locally by default.

Local operation mode: Use the keys on the panel to perform relevant operations.

Remote Operation Mode: The electronic load is connected to a PC, and the PC performs relevant operations on the electronic load. When the electronic load is in remote mode, all panel keys are inactive except for the SHIFT + On/Off key. The mode can be switched to Local Operation Mode by long-pressing the SHIFT + On/Off key. Changing the constant mode operation will not affect the settings of the electronic load.
Constant Mode Operation

The electronic load can operate in the following four constant modes:

Constant Current mode (CC)

Constant Voltage mode (CV)

Constant Resistance mode (CR)

Constant Power mode(CP)

Constant Current Mode(CC)
When the user selects the CC function key, the load enters constant current mode. In constant current operation mode, regardless of changes in the input voltage, the electronic load consumes a constant current. The relationship between voltage and current is shown in the figure below.
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	Figure 4.3.1-1 CC mode: The Voltage-Current Relationship Diagram
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	Figure 4.3.1-2 CC mode operation interface


There are two ways to set the output current (CC mode):

Mode one:

Press MODE and rotary the knob to select CC mode.

Press ENTER key to edit current setting, turn the knob to directly adjust to the set value. During the process of inputting a value, pressing the knob will increment the value to the right. Once it reaches the highest digit, pressing it again will return to the lowest digit, and this will cycle continuously.
Press On/Off key to output the setting current.

Mode two：

Press CC , at the same time, the user interface will switch into CC mode.

Press ENTER key to edit current setting, turn the knob to directly adjust to the set value. During the process of inputting a value, pressing the knob will increment the value to the right. Once it reaches the highest digit, pressing it again will return to the lowest digit, and this will cycle continuously.
Press On/Off key to output the setting current.

In constant current mode, the user can also set the following parameters.

Operation steps:

Press F1 to enter parameter setting interface

Set the maximum working current

Set voltage/current range

Set the rising slope of current

Set the falling slope of current

OCP, OVP, OPP setting

Constant Voltage Mode(CV)
When the user selects the CV function key, the load enters constant voltage mode. In constant voltage operation mode, the electronic load will draw sufficient current to maintain the input voltage at the set value. The relationship between voltage and current is shown in the figure below.
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	Figure 4.3.2-1 CV mode: The Voltage-Current Relationship Diagram

	[image: image20.png])
CVSETUP

OFF CV

=B =
e e e
B 55
EY £3 &5
e e
[ i i

T 00:00:00

MODE

PARAM.

MENU





	Figure 4.3.2-2 CV mode operation interface


There are two ways to set the output voltage (CV mode):

Mode one:

Press MODE and rotary the knob to select CV mode.

Press Enter key to edit voltage setting, turn the knob to directly adjust to the set value. During the process of inputting a value, pressing the knob will increment the value to the right. Once it reaches the highest digit, pressing it again will return to the lowest digit, and this will cycle continuously.
Press On/Off key to output the setting voltage.

Mode two:

Press CV , at the same time, the user interface will switch into CV mode.

Press Enter key to edit voltage setting, turn the knob to directly adjust to the set value. During the process of inputting a value, pressing the knob will increment the value to the right. Once it reaches the highest digit, pressing it again will return to the lowest digit, and this will cycle continuously.
Press On/Off key to output the setting voltage.

In constant voltage mode, the user can also set the following parameters.

Operation steps:

Press F1 to enter parameter setting interface

Set the maximum working current

Set the voltage range

Set the rising/falling slope of the voltage

Fast b) Normal c) Slow

OCP, OVP, OPP setting

Constant Resistance Mode (CR)

When the user selects the SHIFT+ CV function key, the load enters constant resistance mode (CR). In constant resistance operation mode, the electronic load will behave as a constant resistor. The current will change linearly with the input voltage. The relationship between voltage and current is shown in the figure below.

	
[image: image21.emf]Load 

voltage

Load current

CR mode

The slope of the 

straight line is 

equal to the set 

resistance value 



	Figure 4.3.3-1 CR mode: The Voltage-Current Relationship Diagram
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	Figure 4.3.3-2 CR mode operation interface


There are two ways to set the output resistance (CR mode):
Mode one:

Press MODE and rotary the knob to select CR mode.

Press Enter key to edit resistance setting, turn the knob to directly adjust to the set value. During the process of inputting a value, pressing the knob will increment the value to the right. Once it reaches the highest digit, pressing it again will return to the lowest digit, and this will cycle continuously.
Press On/Off key to output the setting resistance.

Mode two:

Press SHIFT+CV , at the same time, the user interface will switch into CR mode.

Press Enter key to edit resistance setting, turn the knob to directly adjust to the set value. During the process of inputting a value, pressing the knob will increment the value to the right. Once it reaches the highest digit, pressing it again will return to the lowest digit, and this will cycle continuously.
Press On/Off key to output the setting resistance.

In constant resistance mode, the user can also set the following parameters.

Operation steps

Press F1 to enter parameter setting interface

Set the maximum working current

Set voltage range

Set the rising/falling slope of resistance

Fast b) Normal c) Slow

OCP, OVP, OPP setting

Constant Power Mode (CP)

When the user selects the SHIFT+ CC function key, the load enters constant power mode. The electronic load will consume a constant power, as shown in the figure below. If the input voltage increases, the input current will decrease. The relationship between voltage and current is shown in the figure below.
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	Figure 4.3.4-1 CP mode: The Voltage-Current Relationship Diagram
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	Figure 4.3.4-2 CP mode operation interface


There are two ways to set the output power (CP mode):
Mode one:

Press MODE and rotary the knob to select CP mode.

Press Enter key to edit power setting, turn the knob to directly adjust to the set value. During the process of inputting a value, pressing the knob will increment the value to the right. Once it reaches the highest digit, pressing it again will return to the lowest digit, and this will cycle continuously.
Press On/Off key to output the setting power.

Mode two:

Press SHIFT+CC , at the same time, the user interface will switch into CP mode.

Press Enter key to edit power setting, turn the knob to directly adjust to the set value. During the process of inputting a value, pressing the knob will increment the value to the right. Once it reaches the highest digit, pressing it again will return to the lowest digit, and this will cycle continuously.
Press On/Off key to output the setting power.

In constant power mode, the user can also set the following parameters.

Operation steps

Press F1 to enter parameter setting interface

Set the maximum working current

Set voltage range

Set the rising/falling slope of power

Fast b) Normal c) Slow

OCP, OVP, OPP setting

Input Control function

After the electronic load is turned on, it is in an unconnected (no load) state. The input switch of the electronic load can be controlled by pressing the On/Off button on the front panel of the electronic load.

When the On/Off indicator light is on,it means the input is turned on,and the TFT screen displays the current real-time voltage and current information.

When the On/Off indicator light is off, the screen will display OFF, 

indicating that the input is turned off.

Keyboard Lock Function
The front panel buttons of the electronic load can be locked by pressing the SHIFT+On/Off button. The screen will display the LOCK status. In this mode, all buttons are disabled except for unlocking the function by pressing and holding the SHIFT+On/Off button.

USB Screenshot Function
When a USB flash drive is inserted into the front panel, the electronic load will automatically recognize it and display a USB icon in the top right corner of the screen, indicating that the USB drive has been successfully connected.

To take a screenshot, long press knob button(around 2 seconds). The device will save the current screen image to the USB drive. Once the screenshot is successfully saved, the message "Screenshot save done" will be displayed. 

Due to the large file size of the screenshots, they can only be saved to a USB flash drive. After capturing the screenshot, you can view the saved image on a computer by accessing the USB drive.
Short Circuit Simulation Function
The electronic load can simulate a short circuit at the input terminal to test whether the protection function of the device under test (DUT) operates correctly when a short circuit occurs at its output terminal. 

To switch to the short circuit state, press (SHIFT) + 1 (SHORT) on the front panel of the electronic load. The short circuit operation does not affect the current settings, and pressing (SHIFT) + 1 (SHORT) or MODE again will return the electronic load to its original set state. The actual current consumed by the electronic load during a short circuit depends on the current operating mode and current range. In CC , CP, and CR modes, when the current is within the maximum range, the short-circuit current will be equal to the current input value. If the current exceeds the maximum range, the maximum short-circuit current will be 110% of the current range. In CP mode, a short circuit is equivalent to setting the load’s constant voltage value to 0V.
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	Figure 4.6-1 Short circuit simulation interface


System Menu Function

Press F1 to enter system setting function.

Detail operation steps:

Press F1 to enter “Function menu” item setting.

Rotary the knob to select the item you want to set.
Press Enter to confirm.
Press Esc or F1 to return the function menu.
	Description:

	To cancel the setting, press the Esc key to return to the function menu     page.

	If you are on any setting page, pressing the CC key or CV key or On/Off will quickly return you to main page.


System Setting

The menu of the electronic load is as follows:

	Menu
	Settings
	Description

	SYSTEM SETUP
	LANGUAGE
	Chinese/ENGLISH

	
	COMM PRO
	SCPI or modbus communication setting

	
	BUTTON VOICE
	Enable or disable button sound

	
	ALARM VOICE
	Enable or disable  alarm sound

	
	BRIGHTNESS
	Screen brightness adjustment from 10% to 100%, with a 10% adjustment each time.

	
	BAUD
	The baud rate setting range of the communication device is from 115200 to 1200; 8 baud rates are available for selection.

	
	RESET
	Restore the settings on the panel to factory defaults.

	
	ADDRESS
	The parallel machine address can be selected from 1 to 10 to set the local machine's parallel address.
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	Figure 4-7-1 system setup page


Parameter Setting

Select System Settings, then choose Parameter Settings. Press the Enter key to enter the Parameter Settings options, as shown in the table below.
The parameter settings of the electronic load are as follows:

	Menu
	Settings
	Description

	Parameter setup
	OVP
	If the input voltage exceeds the set safety threshold (OVP point), the OVP protection will be triggered.

	
	OCP
	If the input current exceeds the set safety threshold (OCP point), the OCP protection will be triggered.

	
	OPP
	If the input power exceeds the set safety threshold (OPP point), the OPP protection will be triggered.

	
	RISE SLOPE
	The setting value for the rise rate of input voltage or current during load operation.

	
	FALL SLOPE
	The setting value for the fall rate of input voltage or current during load operation.

	
	CURRENT LIMIT
	The maximum current value that can be set for normal load operation.

	
	REMOTE COMPENSATION
	This option allows enabling the remote sensing function at the rear terminals, with the default setting for compensation at the front panel input terminals.

	
	VOLTAGE RANGE
	The input voltage range supported by the electronic load is as follows:  

Low range: 0-15V  

High range: 0-150V

	
	LOAD VOLTAGE
	After the load is connected, normal load operation can occur once the input voltage reaches the set load voltage.
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	Figure 4-7-2 parameters setup page


Protection Functions and Fault Information

Products monitor the internal operation of the machine in real time, including system logic faults, external mains power faults, internal power stage faults such as OVP (Over Voltage Protection), OCP (Over Current Protection), OPP (Over Power Protection), OTP (Over Temperature Protection), as well as parallel operation faults, input reverse voltage protection, etc.

When any of the above protections are triggered, the electronic load enters the corresponding self-protection mode. For example, in the case of over-temperature protection, the load will trigger an alarm and automatically switch to the OFF state, with the main unit displaying OTP.

Protection mode can be cleared by pressing the power button or the ENTER button on the front panel.

Over Voltage Protection (OVP)
If the over-voltage circuit is triggered, after a 100us delay, the electronic load will immediately turn OFF, and the buzzer will sound. The main unit will display OVP. Protection can be cleared by pressing the power button or the ENTER button.

Over Current Protection (OCP)

There are two types of over-current protection in the electronic load:

Hardware overcurrent protection: The maximum load current of the instrument will be limited to the corresponding current value by the hardware maximum power (such as OEL1512: 400W). Once the hardware overcurrent protection is triggered, the machine is at the maximum power value and will not trigger the current On/Off state.

Software Over Current Protection: When the software overcurrent protection function is turned on, if the load current exceeds the overcurrent protection setting value, after 100ms, the load will automatically turn off, and the panel will display OCP. In order to ensure the reliability of the machine, if the load current abnormally exceeds 2 times the rated maximum current of the machine for 40us or 1.1 times the rated current for 100us, the fast software overcurrent protection is triggered, the load will automatically turn off, and the panel will display OCP. Users can set the software overcurrent protection value of the  product. For related configuration, please refer to F1 key to enter the parameter setting.

Over Power Protection (OPP)

There are two types of over-power protection in the electronic load:

Hardware Over Power Protection: The maximum load power will be limited to the corresponding power value by the hardware maximum power (such as OEL1512: 400W). Once the hardware over-power protection is triggered, the machine is at the maximum power value and will not trigger the current On/Off state.

Software Over Power Protection: When the software over-power protection function is turned on, if the load power exceeds the over-power protection setting value and lasts for 100ms, the load will automatically turn off and the panel will display OPP. In order to ensure the reliability of the machine, if the load power abnormally exceeds 2 times the rated maximum current of the machine for 40us or 1.1 times the rated current for 100us, the fast software over-power protection is triggered, the load will automatically turn off and the panel will display OPP. Users can set the software over-power protection value of the load. For related configuration items, please refer to F1 to enter the parameter setting.

Over Temperature Protection (OTP)

When the internal devices’ temperature of the electronic load exceeds about 85°C, the electronic load would trigger the temperature protection mode, and the electronic load automatically turns OFF and the panel displays OTP.

Local Reverse Voltage Protection (LRV)
When the electronic load detects a reverse voltage input on the front panel, it will trigger the Local Reverse Voltage (LRV) protection warning. In this case, please confirm that the front panel input terminals are correctly connected to the device under test. It can be released by pressing the power button or the ENTER button.

	Fault Type
	ERROR CODE
	Fault cause Description

	MAINS AC FAULT
	0x0002
	Abnormal mains undervoltage

	INTERNAL FAULT
	0x0004
	Over temperature protection

	
	0x0010
	Error in internal parameter storage

	
	0x0020
	Model firmware match error

	OUTPUT FAULTS
	0x0002
	Output over voltage protection

	
	0x0008
	Output over current protection

	
	0x0010
	Output over power protection

	
	0x0004
	Local Reverse Voltage Protection

	PARALLEL

FAULTS
	0x0001
	Parallel address conflict

	
	0x0002
	Multiple master fault

	
	0x0008
	Communication drop

	
	0x0010
	Mismatch in operable unit fault


Communication Port Introduction

The rear panel of the load in this series is equipped with a USB serial port. After connecting to a PC, ensure that the serial port settings in the PC software interface match those of the load device (9600 baud rate, 8 data bits, no parity, 1 stop bit).

USB port
Use a USB cable with USB A-type on one end and USB B-type on the other end to connect the load to the computer. All electronic load functions can be controlled through USB programming. After connecting the load and PC via USB, the CH340 USB driver needs to be installed (available for download from the our website or by directly contacting our technical support), otherwise, communication will not be successful.

The electronic load supports RS232 and RS485 (for internal parallel use). The following table provides the corresponding pin definitions (output via DB9 terminal connector): 

	Pin
	Function
	Functions
	Notes

	Pin 2
	RS232
	TXD
	Protocol supports: SCPI, Modbus RTU Typical 3.3V

	Pin 3
	
	RXD
	

	Pin 5
	
	GND
	

	Pin 4
	External trigger for transient command
	TRIG IN+
	3.3V TTL signal

	Pin 9
	
	TRIG IN-
	

	Pin 1
	External trigger out for power on/off signal indicator
	TRIGOUT+
	Floating, needed 1k resistor pull up to an external power rail 5V or 3.3V

	Pin 8
	
	TRIGOUT-
	

	Pin 6
	RS485
	RS485_A
	Optional; please make sure there is no signal in these pins when using RS232.

	Pin 7
	
	RS485_B
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Dynamic Mode Function

In dynamic mode, the electronic load allows the user to set two load current parameters corresponding to the constant mode. By configuring the operation mode, the load current will switch back and forth between the two set values.
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Dynamic Mode Setting Parameters:
	Dynamic Test
	Description

	Operation Mode
	Select the continuous operation mode: Continuous / Single

	T1 Value
	Set the T1 load current parameter

	T1 Time
	Set the T1 load current duration

	T2 Value
	Set the T2 load current parameter

	T2 Time
	Set the T2 load current duration

	Rise Slope
	Set the rise slope

	Fall Slope
	Set the fall slope


Continuous Mode

In continuous mode, when dynamic test operation is enabled, the load will continuously switch between the T1 value and T2 value.
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Single Pulse Mode (Pulse)

In single pulse mode, when dynamic test operation is enabled, the load will switch once between the T1 value and T2 value. Each time the load receives a “SHIF+.” trigger command, it will perform a single switching action. The diagram below shows a schematic of the operation interface in this mode.
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VON Function
When testing some power products with slow voltage rise speed, if you turn on the input of the electronic load first and then turn on the power supply, the power supply may be pulled to protect. To this end, the user can set the Von value. When the power supply voltage is higher than this value, the electronic load will start to pull the load. For this series of electronic load products, when the power supply voltage to be tested drops and is less than the Von Point unloading voltage, the load will not unload.
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Note: Please confirm whether you need to set the load voltage. Setting the load voltage is to facilitate users to limit the working voltage value. If it does not need to be limited, please do not set it arbitrarily to avoid the trouble of not being able to load. If the instrument cannot load, please first check whether the VON function is set to a value other than 0. If it is set, please reset the Von value to the minimum value of 0.

Compensation Function

When the electronic load operates with high current, significant voltage drops can occur across the connection cables between the load and the device under test (DUT). To ensure measurement accuracy, the electronic load provides a pair of remote compensation terminals. Users can use these terminals to sample the precise output voltage of the DUT.

Operating Steps:
In the system parameter settings page, locate the remote compensation (RComp) option and select "Rear." Press confirm to enable the remote compensation function. At this point, the remote sampling voltage will be used as the actual detected voltage.

Return to the parameter settings page and select the remote compensation (RComp) option as "Front." This will disable the remote compensation function for the rear terminal, and the load will measure the voltage at the main front-panel binding posts as the actual voltage.

Note:

After enabling rear remote compensation, ensure correct polarity connections between the DUT and the terminals. Do not leave connections floating or reverse the polarity! Do not switch the sense mode while the system is running！
Test Function

Battery discharge test function (Only available for T suffix version)

The electronic load can use the Battery mode to simulate the battery discharge test. After selecting the discharge test mode, you can set the test mode termination conditions: "cut-off voltage (Vstop)", "battery capacity (Cstop)" and "discharge time (Tstop)". When any of the three conditions is met, the discharge stops and the electronic load automatically switches to the OFF state.

When the discharge test uses only one or two of the conditions as the termination judgment conditions, the other unused conditions must be set to "OFF".

Press the On/Off button to turn on the channel. During the test, the main interface can display the actual current, battery discharge voltage, battery discharge time, and battery discharged capacity in real time. This test can reflect the reliability and remaining life of the battery, so it is very necessary to perform this test before replacing the battery.

If the battery discharge termination conditions are met, after the test is completed, the load interface pops up a "Battery discharge test completed!" prompt box, and the user presses the Enter key to exit the prompt box.
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Tip: If the cut-off voltage is used as the termination condition, when the battery voltage is too low (lower than the cut-off voltage), the system determines that the battery has reached the set value, the machine will consider the test completed and automatically interrupt the test. If the cut-off voltage condition is OFF, when the port voltage is detected to be 0V, the electronic load will report a discharge abnormality, automatically terminate the discharge, and the machine will shut down.

During the discharge process, if you need to pause or terminate the battery discharge test, you can select the corresponding option through the On/Off key in the prompt box that appears on the load interface: Pause Discharge, Terminate Discharge. After confirming the selection, enter OFF, and the key indicator will turn off.
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If the discharge is terminated, press the On/Off key again to power on, starting a new round of discharge recording.
To exit the battery test mode, select Programming Mode from the system menu and change the Battery Mode to General Mode.
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Note: To avoid electric shock, please ensure that the battery and electronic load terminals are properly connected before turning on the channel input.

Battery Testing Steps:

In Programming Mode, select Battery Mode, set the CC Current, and the charging stop conditions (Vstop, Cstop, Tstop).

Press the On/Off key to start the test. The panel will display the discharge voltage, current, discharge time, and capacity.

Press On/Off to pause or terminate the test. You can exit the battery capacity test mode by selecting Menu/Programming Mode/General Mode.

Battery Connection Precautions: When using a switch-type device to directly connect the device under test and the electronic load, voltage overshoot may occur at the moment the switch closes. Excessive voltage overshoot can damage the electronic load. To avoid equipment damage, it is recommended to use the following wiring method in battery discharge testing to ensure the safe operation of the test equipment.
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NFB: Non-fuse breaker, used to immediately cut off the circuit in case of an internal aging short circuit in the load. The current rating of the non-fuse breaker must be less than the maximum current rating of the battery.
R: Current-limiting resistor, used to prevent excessive surge current in the battery when the NFB1 is activated, due to the RC of the load port. This helps protect the battery under test from damage. Additionally, it can prevent resonance between the parasitic inductance on the test leads and the load port RC when the relay closes, which may cause excessive voltage at the load end, thus avoiding damage to the electronic load. It is recommended that the resistance value of the current-limiting resistor be greater than 100kΩ.
Fuse: A fuse that melts and cuts off the current when the current increases to an abnormal level and the temperature becomes too high, thus protecting the circuit.
Starting the Discharge Test: First, connect the electronic load and the battery as shown in the diagram. Then, close switch NFB1, wait for a 5-second delay, and then close switch NFB2 to begin the battery discharge test.

Stopping the Discharge Test: First, press the [On/Off] button on the front panel of the load to turn off the load input. Then, open switch NFB2, and finally, open switch NFB1 to stop the battery discharge test.

OCP test function (Only available for T suffix version)

This series of electronic loads has an overcurrent protection (OCP) test function. In the OCP test mode, when the input voltage reaches the Von value, a delay period is set, and the electronic load works. The current increases by the step value at regular intervals, and the load input voltage is detected at the same time to determine whether it is higher than the OCP voltage value. If it is higher, it indicates that OCP has not occurred, and the current step operation is repeated until it reaches the cut-off current; if it is lower, it indicates that OCP has occurred, and then check whether the current value is within the target range. If it is within the range, it is Pass, otherwise it is Fault.
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OCP test steps:

1. Select OCP mode in programming mode, enter the OCP test function setting interface, and edit the OCP setting parameters.

2. Press the On/Off button to start the OCP test. If it is within the range, the test is PASS and the panel displays the corresponding display. If it is out of the range, the test fails and the panel displays the corresponding display.

3. After the test is completed, the user needs to return to the settings interface and reset it.
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OPP Test Function (Only available for T suffix version)
This series of electronic loads has an over-power protection (OPP) test function. In the OPP test mode, when the input voltage reaches the Von value, the load starts to work after a delay. The current increases by the step value at regular intervals, and the load input voltage is detected at the same time to determine whether it is higher than the OPP voltage value. If it is higher, it means that OPP has not occurred, and the power step operation is repeated until it reaches the cut-off power; if it is lower, it means that OPP has occurred, and then check whether the current power value is within the target range. If it is within the range, it is Pass, otherwise it is Fault.
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OPP test steps:

1. Select OPP test mode in programming mode, enter the OPP test function setting interface, and edit the OPP setting parameters.

2. Press the On/Off button to start the OPP test. If it is within the range, the test is PASS and the panel displays the corresponding display. If it is out of the range, the test fails and the panel displays the corresponding display.
3. After the test is completed, the user needs to return to the settings interface and reset it.
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Ripple Test Function (Only available for T suffix version)
This series of electronic loads is optionally equipped with a ripple test function, which can well meet the ripple measurement demands of switching power supplies and chargers. Users can directly connect the power supply to the load without using an additional oscilloscope, and read voltage ripple values (Vpp / Vpp+ / Vpp-) and current ripple values (Ipp / Ipp+ / Ipp-) directly. It greatly simplifies experimental wiring and operation. The system supports ripple display, screen capture and automatic statistical analysis. It is highly suitable for fast and batch ripple inspection scenarios, especially for R&D initial debugging and production line quality inspection.

The ripple sampling rate can be set to five options: 1kHz, 5kHz, 10kHz, 25kHz and 50kHz, with the display time ranging from 16ms to 800ms. When entering the ripple interface, if the user is unsure of the ripple frequency of the device under test, select the lowest sampling rate of 1kHz to display a complete AC cycle on the entire screen. Then adjust the sampling frequency to 10kHz or 25kHz according to the ripple density to view the actual ripple waveform in detail.
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Ripple test operations:
On the main page, press SHIFT+F2 to quickly enter the ripple test function interface.
Use the knob to select ripple voltage or ripple current to be tested. After confirmation, set the sampling rate. For unknown ripple signals, first select the lowest sampling rate to display a complete ripple cycle on the screen, then zoom in for detailed observation.
After setting, press the RUN key on the interface to start the ripple test. The ripple waveform as well as voltage and current ripple values will be displayed during the test. To pause and check the ripple magnitude, press the RUN key again to enter the HOLD state, and the screen will stop refreshing.
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PC software introduction

Please go to the our official website or contact your local dealer or our sales specialist to download or request the PC software for OEL15 series Electronic Load

Connect the USB CDC port on the back panel to the host computer. After the instrument is successfully connected, enter the mode setting to remotely operate the instrument.
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Troubleshooting

The instrument is powered on but no display. Please following the steps:

Check if the power is connected properly.
Check if the fuse which is below the AC Power socket is used appropriately and in good condition (the cover can be pried open with a straight screwdriver).
Restart the instrument after the steps above.
If the problem still exists, please contact our customer service.
Technical Specification
The instrument must be operated continuously for more than 30 minutes at the specified operating temperature to achieve the following specifications:
	Model
	OEL

1530
	OEL

1520
	OEL

1515
	OEL

3030
	OEL

3020
	OEL

3015
	OEL

6030
	OEL

6020
	OEL

6015

	Input
	Voltage
	0-150V
	0-300V
	0-600V

	
	Current
	0-30A
	0-20A
	0-15A
	0-30A
	0-20A
	0-15A
	0-30A
	0-20A
	0-15A

	
	Power
	300W
	150W
	300W
	150W
	300W
	150W

	Minimum Operation voltage
	1.1V@30A
	0.9V@20A
	0.8V@15A
	1.1V@30A
	0.9V@20A
	0.8V@15A
	1.1V@30A
	0.9V@20A
	0.8V@15A

	Constant voltage mode (CV)

	Range
	0-150V
	0-300V
	0-600V

	Resolution
	1mV

	Accuracy
	0.05%+0.05%FS
	0.05%+0.05%FS
	0.05%+0.05%FS

	Constant current mode (CC)

	Range
	0-30A
	0-20A
	0-15A
	0-30A
	0-20A
	0-15A
	0-30A
	0-20A
	0-15A

	Resolution
	1mA
	1mA
	1mA

	Accuracy
	0.05%+0.05%IFS
	0.05%+0.05%IFS
	0.05%+0.05%IFS

	Constant resistance mode (CR)

	Range
	0.05-7.5k Ω
	0.05-7.5k Ω
	0.05-7.5k Ω

	Accuracy*1
	0.1%+0.01%R
	0.1%+0.01%R
	0.1%+0.01%R

	Constant power mode (CP)

	Range
	300W
	150W
	300W
	150W
	300W
	150W

	Resolution
	10mW
	10mW
	10mW

	Accuracy*1
	0.1%+0.1%FS
	0.1%+0.1%FS
	0.1%+0.1%FS

	Dynamic mode (CCD)

	T1&T2
	0.1ms-50s
	0.1ms-50s
	0.1ms-50s

	Rise/Fall slope
	0.01A/ms-2000A/ms
	0.01A/ms-2000A/ms
	0.01A/ms-2000A/ms

	Min rise time
	20us
	20us
	20us

	Voltage readback value

	Range
	0-15V/
0-150V
	0-30V/
0-300V
	0-60V/
0-600V

	Resolution
	1mV
	1mV
	1mV

	Accuracy
	0.05%+0.05%FS
	0.05%+0.05%FS
	0.05%+0.05%FS

	Current readback value

	Range
	0-30A
	0-20A
	0-15A
	0-30A
	0-20A
	0-15A
	0-30A
	0-20A
	0-15A

	Resolution
	0.1mA
	0.1mA
	0.1mA

	Accuracy
	0.05%+0.05%FS
	0.05%+0.05%FS
	0.05%+0.05%FS

	Power readback value

	Range
	300W
	150W
	300W
	150W
	300W
	150W

	Resolution
	10mW
	10mW
	10mW

	Accuracy
	0.1%+0.1%FS
	0.1%+0.1%FS
	0.1%+0.1%FS

	Protection function

	OPP
	>OPP value 10ms protection
	>OPP value 10ms protection
	>OPP value 10ms protection

	OVP 
	>OVP value 0.1ms protection
	>OVP value 0.1ms protection
	>OVP value 0.1ms protection

	OCP
	>OCP value 0.1ms protection
	>OCP value 0.1ms protection
	>OCP value 0.1ms protection

	OTP
	85℃
	85℃
	85℃

	Short circuit

	Current (CC)
	>30A
	>20A
	>15A
	>30A
	>20A
	>15A
	>30A
	>20A
	>15A

	Voltage (CV)
	0
	0
	0

	Resistance (CR)
	50mΩ
	50mΩ
	50mΩ

	Test Function

	Test Function
	Battery test, OCP,OPP, Optional, only available for “T”version
Ripple test function, Optional, only available for “V”version

	AC parameter

	Voltage range
	100V-240Vac

	Frequency range
	50/60Hz

	Interface

	RS-232 /RS458
	Support RS232, RS485, USB, Supported SCPI

	General

	Heat dissipation
	Internal intelligent fan forced air cooling

	Withstand voltage (Input to Ground)
	1.5KV/AC/5mA

	Volume (L*W*H)
	282.9mm*84.3mm*151.15mm

	Net weight 
	Approx. 2.5kg
	Approx. 2.5kg
	Approx. 2.5kg


* 1 Accuracy: The voltage and current input values should not be less than 10%FS

If there are any changes to parameters, no separate notice will be given. Please refer to the website update for the latest information.

Interval Period of Adjustment: One year is recommended for the calibration interval period.
Appendix

Appendix A: Accessories

(The accessories subject to final delivery.)

	Standard
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Power Cord
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User Manual
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Fuse
	
	

	Optional
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Remote Compensation Cable

(Order No. 7CL3-10)
	[image: image51.png]


USB Cable (Type A to Type C)

(Order No. 7ULC08)
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RS232

(Order No. 7R232MFS)
	
	


Appendix B: General Care and Cleaning
General Care

Do not store or leave the instrument where the liquid crystal display could be exposed to direct sunlight for long periods of time.
Caution: To avoid any damage to the instrument, do not exposed it to any sprays, liquids, or solvents.
Cleaning
Inspect the instrument as often as operating conditions require. To clean the instrument exterior, perform the following steps:
Wipe the dust from the instrument surface with a soft cloth. Take care not to scratch the transparent LCD protection screen when cleaning.
Disconnect power before cleaning your instrument. Clean the instrument with a damp soft cloth (not dripping with water). It is recommended to clean with soft detergent or fresh water. To avoid damage to the instrument, do not use any corrosive chemical cleaning agents.
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	Warning: 

Before re-applying power, ensure that the instrument is completely dry, avoiding any electric shock or electrical short circuit resulting from moisture.
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