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The instrument must be operated continuously for more than 30 minutes at the specified operating temperature to achieve the following specifications:
	Model
	OEL8512(T)
	OEL8512B
(T)
	OEL8512C
(T)
	OEL8511(T)
	OEL8511B
(T)
	OEL8511C
(T)
	OEL8513
(T)

	Input
	Voltage
	0-150V
	0-300V
	0-600V
	0-150V
	0-300V
	0-600V
	0-150V

	
	Current
	0-40A
	0-20A
	0-40A

	
	Power
	400W
	200W
	500W

	Minimum Operation voltage
	1.1V@40A
	0.8V@20A
	1.1V@40A

	Constant voltage mode (CV)

	Range
	0-150V
	0-300V
	0-600V
	0-150V
	0-300V
	0-600V
	0-150V

	Resolution
	1mV
	1mV
	1mV

	Accuracy
	0.05%+0.05%FS

	Constant current mode (CC)

	Range
	0-40A
	0-20A
	0-40A

	Resolution
	1mA
	1mA
	1mA

	Accuracy
	0.03%+0.03%IFS

	Constant resistance mode (CR)

	Range
	0.05-7.5k 

	Accuracy *1
	0.1%+0.01%R

	Constant power mode (CP)

	Range
	400W
	200W
	500W

	Resolution
	10mW
	10mW
	10mW

	Accuracy *1
	0.1%+0.1%FS

	Dynamic mode (CCD)

	T1&T2
	0.1ms-50s

	Rise/Fall slope
	0.01A/ms-2000A/ms

	Min rise time
	20us

	Voltage readback value

	Range
	0-15V
	0-150V
	0-30V
	0-300V
	0-60V
	0-600V
	0-15V
	0-150V
	0-30V
	0-300V
	0-60V
	0-600V
	0-150V

	Resolution
	1mV
	1mV
	1mV
	1mV
	1mV
	1mV
	1mV

	Accuracy
	0.03%+
0.05%FS
	0.03%+
0.05%FS
	0.03%+
0.05%FS
	0.03%+
0.05%FS
	0.03%+
0.05%FS
	0.03%+
0.05%FS
	0.03%+
0.05%FS

	Current readback value

	Range
	0-40A
	0-20A
	0-40A

	Resolution
	0.1mA
	0.1mA
	0.1mA

	Accuracy
	0.03%+0.03%FS

	Power readback value

	Range
	0-400W
	0-200W
	0-500W

	Resolution
	10mW

	Accuracy
	0.1%+0.1%FS

	Protection function

	OPP
	>OPP value 10ms protection

	OVP
	>OVP value 0.1ms protection

	OCP 
	>OCP value 0.1ms protection

	OTP
	85℃

	Short circuit

	Current (CC)
	>40A
	>20A
	>40A

	Voltage (CV)
	0

	Resistance (CR)
	50mΩ

	Test Function

	Battery test, OCP
	Optional, only for “T”version

	AC parameter

	Voltage range
	110V ±10% or 220V ±10%

	Frequency range
	50/60Hz

	Interface

	RS-232 
	Supported SCPI

	General

	Heat dissipation
	Internal intelligent fan forced air cooling 

	Safety
	Compliant with the European EMC Directive 89/336/EEC, for Class A test and measurement products

	Withstand voltage (Input to Ground)
	1.5KV/AC/5mA

	Volume (L*W*H)
	375mm*215mm*88mm

	Net weight 
	Approx. 5kg
	Approx. 4.2kg
	Approx. 5kg


* 1 Accuracy: The voltage and current input values should not be less than 10%FS
If there are any changes to parameters, no separate notice will be given. Please refer to the website update for the latest information.
Interval Period of Adjustment: One year is recommended for the calibration interval period.
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