OEL85 Series Technical Specifications

The instrument must be operated continuously for more than 30 minutes at the specified operating temperature to achieve the following specifications:

| Model |  | OEL8512(T) |  | OEL8512B (T) |  | OEL8512C (T) |  | OEL8511(T) |  | OEL8511B (T) |  | OEL8511C (T) |  | OEL8513 (T) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Input | Voltage | 0-150V |  | 0-300V |  | 0-600V |  | 0-150V |  | 0-300V |  | 0-600V |  | 0-150V |
|  | Current | 0-40A |  |  |  |  |  | 0-20A |  |  |  |  |  | 0-40A |
|  | Power | 400W |  |  |  |  |  | 200W |  |  |  |  |  | 500W |
| Minimum Operation voltage |  | 1.1V@40A |  |  |  |  |  | 0.8V@20A |  |  |  |  |  | 1.1V@40A |
| Constant voltage mode (CV) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Range |  | 0-150V |  | 0-300V |  | 0-600V |  | 0-150V |  | 0-300V |  | 0-600V |  | 0-150V |
| Resolution |  | 1mV |  |  |  |  |  | 1mV |  |  |  |  |  | 1mV |
| Accuracy |  | 0.05%+0.05%FS |  |  |  |  |  |  |  |  |  |  |  |  |
| Constant current mode (CC) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Range |  | 0-40A |  |  |  |  |  | 0-20A |  |  |  |  |  | 0-40A |
| Resolution |  | 1mA |  |  |  |  |  | 1mA |  |  |  |  |  | 1mA |
| Accuracy |  | 0.03%+0.03%IFS |  |  |  |  |  |  |  |  |  |  |  |  |
| Constant resistance mode (CR) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Range |  | 0.05-7.5k |  |  |  |  |  |  |  |  |  |  |  |  |
| Accuracy *1 |  | 0.1%+0.01%R |  |  |  |  |  |  |  |  |  |  |  |  |
| Constant power mode (CP) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Range |  | 400W |  |  |  |  |  | 200W |  |  |  |  |  | 500W |
| Resolution |  | 10mW |  |  |  |  |  | 10mW |  |  |  |  |  | 10mW |
| Accuracy *1 |  | 0.1%+0.1%FS |  |  |  |  |  |  |  |  |  |  |  |  |
| Dynamic mode (CCD) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| T1&T2 |  | 0.1ms-50s |  |  |  |  |  |  |  |  |  |  |  |  |
| Rise/Fall slope |  | 0.01A/ms-2000A/ms |  |  |  |  |  |  |  |  |  |  |  |  |
| Min rise time |  | 20us |  |  |  |  |  |  |  |  |  |  |  |  |
| Voltage readback value |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Range |  | 0-15V | 0-150V | 0-30V | 0-300V | 0-60V | 0-600V | 0-15V | 0-150V | 0-30V | 0-300V | 0-60V | 0-600V | 0-150V |
| Resolution |  | 1mV |  | 1mV |  | 1mV |  | 1mV |  | 1mV |  | 1mV |  | 1mV |
| Accuracy |  | 0.03%+ 0.05%FS |  | 0.03%+ 0.05%FS |  | 0.03%+ 0.05%FS |  | 0.03%+ 0.05%FS |  | 0.03%+ 0.05%FS |  | 0.03%+ 0.05%FS |  | 0.03%+ 0.05%FS |
| Current readback value |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Range |  | 0-40A |  |  |  |  |  | 0-20A |  |  |  |  |  | 0-40A |
| Resolution |  | 0.1mA |  |  |  |  |  | 0.1mA |  |  |  |  |  | 0.1mA |
| Accuracy |  | 0.03%+0.03%FS |  |  |  |  |  |  |  |  |  |  |  |  |
| Power readback value |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Range |  | 0-400W |  |  |  |  |  | 0-200W |  |  |  |  |  | 0-500W |
| Resolution |  | 10mW |  |  |  |  |  |  |  |  |  |  |  |  |
| Accuracy |  | 0.1%+0.1%FS |  |  |  |  |  |  |  |  |  |  |  |  |
| Protection function |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| OPP |  | >OPP value 10ms protection |  |  |  |  |  |  |  |  |  |  |  |  |
| OVP |  | >OVP value 0.1ms protection |  |  |  |  |  |  |  |  |  |  |  |  |
| OCP |  | >OCP value 0.1ms protection |  |  |  |  |  |  |  |  |  |  |  |  |
| OTP |  | 85℃ |  |  |  |  |  |  |  |  |  |  |  |  |
| Short circuit |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Current (CC) |  | >40A |  |  |  |  |  | >20A |  |  |  |  |  | >40A |
| Voltage (CV) |  | 0 |  |  |  |  |  |  |  |  |  |  |  |  |
| Resistance (CR) |  | 50mΩ |  |  |  |  |  |  |  |  |  |  |  |  |
| Test Function |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Battery test, OCP |  | Optional, only for “T”version |  |  |  |  |  |  |  |  |  |  |  |  |
| AC parameter |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Voltage range |  | 110V ±10% or 220V ±10% |  |  |  |  |  |  |  |  |  |  |  |  |
| Frequency range |  | 50/60Hz |  |  |  |  |  |  |  |  |  |  |  |  |
| Interface |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| RS-232 |  | Supported SCPI |  |  |  |  |  |  |  |  |  |  |  |  |
| General |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Heat dissipation |  | Internal intelligent fan forced air cooling |  |  |  |  |  |  |  |  |  |  |  |  |
| Safety |  | Compliant with the European EMC Directive 89/336/EEC, for Class A test and measurement products |  |  |  |  |  |  |  |  |  |  |  |  |
| Withstand voltage (Input to Ground) |  | 1.5KV/AC/5mA |  |  |  |  |  |  |  |  |  |  |  |  |
| Volume (L*W*H) |  | 375mm*215mm*88mm |  |  |  |  |  |  |  |  |  |  |  |  |
| Net weight |  | Approx. 5kg |  |  |  |  |  | Approx. 4.2kg |  |  |  |  |  | Approx. 5kg |

* 1 Accuracy: The voltage and current input values should not be less than 10%FS

If there are any changes to parameters, no separate notice will be given. Please refer to the website update for the latest information.

Interval Period of Adjustment: One year is recommended for the calibration interval period.

V1.0.0