
OWH65 Series Single-Channel High-Power Programmable DC Power Supply Technical Specifications
The instrument must be continuously operated at the specified operating temperature for more than 30 minutes to meet the following specifications:
OWH6546 (A) Parameters
	Model
	OWH6546
	OWH6546A

	Rated output
 (0℃-40℃)
	Output voltage range
	0 -1200V
	0-1200V

	
	OVP
OVP
	1V-1250V
	1V-1250V

	
	Output current range
	0-30A
	0-20A

	
	OCP
	1A-35A
	1A-25A

	
	Max. power
	18kW
	12kW

	Input voltage
	Volt/freq.
	3P4W, 380Vac+/-10%; 45Hz-65Hz (L1/L2/L3-N)

	Load regulation ratio
	CV
	≤300mV
	≤300mV

	
	CC
	≤15mA
	≤45mA

	Power regulation rate
	CV
	≤300mV
	≤300mV

	
	CC
	≤15mA
	≤45mA

	Setting Resolution
	Voltage
	10mV
	10mV

	
	Current
	1mA
	1mA

	Readback resolution
	Voltage
	10mV
	10mV

	
	Current
	1mA
	1mA

	Setting Accuracy
	Voltage
	≤0.05% ± 500mV
	≤0.05% ± 500mV

	
	Current
	≤0.1% ± 30mA
	≤0.1% ± 90mA

	Readback accuracy
	Voltage
	≤0.05% ± 500mV
	≤0.05% ± 500mV

	
	Current
	≤0.1% ± 30mA
	≤0.1% ± 90mA

	Ripple Noise (*)
	Voltage
	≤2Vp-p
	≤2Vp-p

	
	Current
	≤150mArms
	≤300mArms

	Output temperature drift coefficient
(0℃-40℃)
	Voltage
	100ppm/℃
	100ppm/℃

	
	Current
	200ppm/℃
	200ppm/℃

	Read back temperature drift coefficient
	Voltage
	100ppm/℃
	100ppm/℃

	
	Current
	200ppm/℃
	200ppm/℃

	Dynamic response (10%-90% rated load)
	≤5.0ms
	≤5.0ms

	Operating temperature range
	0-40℃
	0-40℃

	Display
	3.9 "color LCD screen
	3.9 "color LCD screen

	Communication interface
	USB, LAN, RS485
	USB, LAN, RS485

	Dimensions (mm)
	428W * 274.4H * 511.5D
	428W * 274.4H * 511.5D

	Weight
	Approx. 30kg
	Approx. 30kg


*: Tested with a 10µF electrolytic capacitor in parallel with a 0.1µF ceramic capacitor at the output, under a noise bandwidth of 20MHz and a ripple bandwidth of 1MHz.
OWH6522 Parameters
	Model
	OWH6522

	Rated output (0 ℃-40 ℃)
	Voltage
	0 - 60V

	
	OVP
	1 - 65V

	
	Current
	0 - 120A

	
	OCP
	1 - 121A

	
	Power
	3000W

	Power supply voltage
	Volt./Freq.
	198V-240Vac; 45Hz-65Hz

	Load regulation ratio *
	CV
	≤30mV

	
	CC
	≤60mA

	Power regulation rate
	CV
	≤30mV

	
	CC
	≤60mA

	Set Resolution
	Voltage
	10mV

	
	Current
	1mA

	Readback resolution
	Voltage
	10mV

	
	Current
	1mA

	Setting accuracy
	Voltage
	≤0.05% ± 30mV

	
	Current
	≤0.1% ± 120mA

	Readback resolution
	Voltage
	≤0.05% ± 30mV

	Readback accuracy
	Voltage
	≤0.1% ± 120mA

	Ripple/noise
	Voltage
	≤80mVp-p

	
	Current
	≤120mArms

	Output temperature coefficient
(0℃-40℃)
	Voltage
	100ppm/℃

	
	Current
	200ppm/℃

	Read back temperature coefficient
	Voltage
	100ppm/℃

	
	Current
	200ppm/℃

	Transient recovery time
(10%-90% rated load)
	≤5.0ms

	Operating temperature
	0-40℃

	Display
	4.3 inch color LCD display

	Interface
	USB, LAN, RS485

	Dimensions (mm)
	428 (W) X 88.0 (H) X 446.2 (D)

	Weight
	Approx. 13kg


*: The detection point of this indicator is the voltage compensation terminal.
Adjustment interval: The recommended calibration interval is one year.
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