
OWH67 Series Programming DC Power Supply Technical Specifications
The instrument must be operated continuously for more than 30 minutes at the specified operating temperature to achieve the following specifications:

OWH67012 Series

	Model
	OWH67012-80
	OWH67012-150
	OWH67012-300

	Rated Output

 (0℃-40℃)
	Voltage
	0 -80V
	0-150V
	0-300V

	
	OVP
	1-85V
	1-155V
	1-305V

	
	Current
	0-30A
	0-20A
	0-10A

	
	OCP
	1-35A
	1-25A
	1-15A

	
	Power 
	1200W

	
	OPP
	10W - 1300W

	Power supply
	Volt/

freq.
	85V-265Vac; 50Hz-60Hz,10A Max

	Load Regulation

(%of output+offset)
	CV
	≤0.03%+0.03%FS

	
	CC
	≤0.03%+0.05%FS

	Line Regulation (%of output+offset)
	CV
	≤0.03%+0.03%FS

	
	CC
	≤0.03%+0.05%FS

	Setting Resolution
	Voltage
	10mV

	
	Current
	1mA

	Readback Resolution
	Voltage
	1mV

	
	Current
	1mA

	Setting Accuracy
	Voltage
	≤0.05%+0.05%FS

	
	Current
	≤0.05%+0.1%FS

	Readback Accuracy
	Voltage
	≤0.05%+0.05%FS

	
	Current
	≤0.05%+0.1%FS

	Ripple/

Noise(*)
	Voltage
	≤100mVp-p
	≤100mVp-p
	≤250mVp-p

	
	Current
	≤50mArms
	≤50mArms
	≤50mArms

	Output temperature coefficient

(0℃-40℃)
	Voltage
	100ppm/℃

	
	Current
	200ppm/℃

	Readback temperature coefficient
	Voltage
	100ppm/℃

	
	Current
	200ppm/℃


*: Noise bandwidth: 20 MHz. Test with a 10 µF electrolytic capacitor connected in parallel with a 0.1 µF ceramic capacitor at the output.
OWH67018 Series

	Model
	OWH67018-150

	Rated Output

 (0℃-40℃)
	Voltage
	0-150V

	
	OVP
	1-155V

	
	Current
	0-30A

	
	OCP
	1-35A

	
	Power 
	1800W

	
	OPP
	10W - 1900W

	Power supply
	Volt/

freq.
	85V-265Vac; 50Hz-60Hz,10A Max

	Load Regulation

(%of output+offset)
	CV
	≤0.03%+0.03%FS

	
	CC
	≤0.03%+0.05%FS

	Line Regulation (%of output+offset)
	CV
	≤0.03%+0.03%FS

	
	CC
	≤0.03%+0.05%FS

	Setting Resolution
	Voltage
	10mV

	
	Current
	1mA

	Readback Resolution
	Voltage
	1mV

	
	Current
	1mA

	Setting Accuracy
	Voltage
	≤0.05%+0.05%FS

	
	Current
	≤0.05%+0.1%FS

	Readback Accuracy
	Voltage
	≤0.05%+0.05%FS

	
	Current
	≤0.05%+0.1%FS

	Ripple/

Noise(*)
	Voltage
	≤100mVp-p

	
	Current
	≤50mArms

	Output temperature coefficient

(0℃-40℃)
	Voltage
	100ppm/℃

	
	Current
	200ppm/℃

	Readback temperature coefficient
	Voltage
	100ppm/℃

	
	Current
	200ppm/℃


*: Noise bandwidth: 20 MHz. Test with a 10 µF electrolytic capacitor connected in parallel with a 0.1 µF ceramic capacitor at the output.
OWH67020 Series

	Model
	OWH67020-150
	OWH67020-300
	OWH67020-600

	Rated Output

 (0℃-40℃)
	Voltage
	0-150V
	0-300V
	0-600V

	
	OVP
	1-155V
	1-305V
	1-605V

	
	Current
	0-20A
	0-10A
	0-10A

	
	OCP
	1-25A
	1-15A
	1-11A

	
	Power 
	2000W

	
	OPP
	10W-2100W

	Power supply
	Volt/

freq.
	85V-265Vac; 50Hz-60Hz,10A Max

	Load Regulation

(%of output+offset)
	CV
	≤0.03%+0.03%FS

	
	CC
	≤0.03%+0.05%FS

	Line Regulation (%of output+offset)
	CV
	≤0.03%+0.03%FS

	
	CC
	≤0.03%+0.05%FS

	Setting Resolution
	Voltage
	10mV

	
	Current
	1mA

	Readback Resolution
	Voltage
	1mV

	
	Current
	1mA

	Setting Accuracy
	Voltage
	≤0.05%+0.05%FS

	
	Current
	≤0.05%+0.1%FS

	Readback Accuracy
	Voltage
	≤0.05%+0.05%FS

	
	Current
	≤0.05%+0.1%FS

	Ripple/

Noise(*)
	Voltage
	≤100mVp-p
	≤250mVp-p
	≤500mVp-p

	
	Current
	≤30mArms
	≤50mArms
	≤50mArms

	Output temperature coefficient

(0℃-40℃)
	Voltage
	100ppm/℃

	
	Current
	200ppm/℃

	Readback temperature coefficient
	Voltage
	100ppm/℃

	
	Current
	200ppm/℃


*: Noise bandwidth: 20 MHz. Test with a 10 µF electrolytic capacitor connected in parallel with a 0.1 µF ceramic capacitor at the output.
OWH67030 Series

	Model
	OWH67030-150
	OWH67030-150A
	OWH67030-300
	OWH67030-600

	Rated Output

 (0℃-40℃)
	Voltage
	0-150V
	0-150V
	0-300V
	0-600V

	
	OVP
	1-155V
	1-155V
	1-305V
	1-605V

	
	Current
	0-30A
	0-40A
	0-15A
	0-10A

	
	OCP
	1-35A
	1-45A
	1-20A
	1-11A

	
	Power 
	3000W

	
	OPP
	10W - 3100W

	Power supply
	Volt/

freq.
	85V-265Vac; 50Hz-60Hz,16A Max

	Load Regulation

(%of output+offset)
	CV
	≤0.03%+0.03%FS

	
	CC
	≤0.03%+0.05%FS

	Line Regulation (%of output+offset)
	CV
	≤0.03%+0.03%FS

	
	CC
	≤0.03%+0.05%FS

	Setting Resolution
	Voltage
	10mV

	
	Current
	1mA

	Readback Resolution
	Voltage
	1mV

	
	Current
	1mA

	Setting Accuracy
	Voltage
	≤0.05%+0.05%FS

	
	Current
	≤0.05%+0.1%FS

	Readback Accuracy
	Voltage
	≤0.05%+0.05%FS

	
	Current
	≤0.05%+0.1%FS

	Ripple/

Noise(*)
	Voltage
	≤100mVp-p
	≤150mVp-p
	≤250mVp-p
	≤500mVp-p

	
	Current
	≤50mArms
	≤50mArms
	≤50mArms
	≤50mArms

	Output temperature coefficient

(0℃-40℃)
	Voltage
	100ppm/℃

	
	Current
	200ppm/℃

	Readback temperature coefficient
	Voltage
	100ppm/℃

	
	Current
	200ppm/℃


*: Noise bandwidth: 20 MHz. Test with a 10 µF electrolytic capacitor connected in parallel with a 0.1 µF ceramic capacitor at the output.
OWH67060 Series

	Model
	OWH67060-300
	OWH67060-600
	OWH67060-1200

	Rated Output

 (0℃-40℃)
	Voltage
	0-300V
	0-600V
	0-1200V

	
	OVP
	1-305V
	0-605V
	0-1250V

	
	Current
	0-30A
	0-20A
	0-10A

	
	OCP
	1-35A
	0-25A
	0-11A

	
	Power 
	6000W

	
	OPP
	10W - 6500W

	Power supply
	Volt/

freq.
	100V-240Vac; 50Hz-60Hz,32A Max

	Load Regulation

(%of output+offset)
	CV
	≤0.03%+0.03%FS

	
	CC
	≤0.03%+0.05%FS

	Line Regulation (%of output+offset)
	CV
	≤0.03%+0.03%FS

	
	CC
	≤0.03%+0.05%FS

	Setting Resolution
	Voltage
	10mV

	
	Current
	1mA

	Readback Resolution
	Voltage
	1mV

	
	Current
	1mA

	Setting Accuracy
	Voltage
	≤0.05%+0.05%FS

	
	Current
	≤0.05%+0.1%FS

	Readback Accuracy
	Voltage
	≤0.05%+0.05%FS

	
	Current
	≤0.05%+0.1%FS

	Ripple/

Noise(*)
	Voltage
	≤500mVp-p
	≤500mVp-p
	≤2Vp-p

	
	Current
	≤50mArms
	≤50mArms
	≤50mArms

	Output temperature coefficient

(0℃-40℃)
	Voltage
	100ppm/℃

	
	Current
	200ppm/℃

	Readback temperature coefficient
	Voltage
	100ppm/℃

	
	Current
	200ppm/℃


*: Noise bandwidth: 20 MHz. Test with a 10 µF electrolytic capacitor connected in parallel with a 0.1 µF ceramic capacitor at the output.
General Specifications

	Characteristics
	Description

	Recovery Time

(10%-90% rated load)
	5ms

	Working Temperature
	0-40℃

	Display
	3.9 inch color LCD display

	Interface
	USB, RS485, LAN(Optional)

	Size
	OWH67012 Series

OWH67018 Series

OWH67020 Series

OWH67030 Series
	214.6mm (W)X 88.0mm (H) X453.0mm(D)

	
	OWH67060 Series
	430.0mm(W) X 88.0mm (H) X453.0mm(D)

	Weight
	OWH67012 Series
	Approx. 6.4kg

	
	OWH67018 Series OWH67020 Series

OWH67030 Series
	Approx. 6.8kg

	
	OWH67060 Series
	Approx. 13kg


Interval Period of Adjustment: 

One year is recommended for the calibration interval period.
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