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Fujian LILLIPUT Optoelectronics Technology Co.,Ltd. 
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Tel: +86-596-2130430
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General Warranty

We warrant that the product will be free from defects in materials and workmanship for a period of 2 years from the date of purchase of the product by the original purchaser from our company. The warranty period for accessories such as probes, battery is 12 months. This warranty only applies to the original purchaser and is not transferable to a third party. 

If the product proves defective during the warranty period, we will either repair the defective product without charge for parts and labour, or will provide a replacement in exchange for the defective product. Parts, modules and replacement products used by our company for warranty work may be new or reconditioned like new. All replaced parts, modules and products become the property of our company.

In order to obtain service under this warranty, the customer must notify our company of the defect before the expiration of the warranty period. Customer shall be responsible for packaging and shipping the defective product to the designated service centre, a copy of the customers proof of purchase is also required.

This warranty shall not apply to any defect, failure or damage caused by improper use or improper or inadequate maintenance and care. We shall not be obligated to furnish service under this warranty a) to repair damage resulting from attempts by personnel other than our company representatives to install, repair or service the product; b) to repair damage resulting from improper use or connection to incompatible equipment; c) to repair any damage or malfunction caused by the use of not our supplies; or d) to service a product that has been modified or integrated with other products when the effect of such modification or integration increases the time or difficulty of servicing the product. 

Please contact the nearest Sales and Service Offices for services.

Excepting the after-sales services provided in this summary or the applicable warranty statements, we will not offer any guarantee for maintenance definitely declared or hinted, including but not limited to the implied guarantee for marketability and special-purpose acceptability. We should not take any responsibilities for any indirect, special or consequent damages.
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General Safety Requirement
Before any operations, please read the following safety precautions to avoid any possible bodily injury and prevent this product or any other products connected from damage. In order to avoid any contingent danger, this product is only used within the range specified.
Only the qualified technicians can implement the maintenance.
To avoid Fire or Personal Injury: 

Use Proper Power Cord. Use only the power cord supplied with the product and certified to use in your country. 

Product Grounded. This instrument is grounded through the power cord grounding conductor. To avoid electric shock, the grounding conductor must be grounded. The product must be grounded properly before any connection with its input or output terminal. 

Limit operation to the specified measurement category, voltage, or amperage ratings.

Check all Terminal Ratings. To avoid fire or shock hazard, check all ratings and markers on the instrument. Refer to the user's manual for more information about ratings before connecting the instrument. Do not exceed any of ratings defined in the following section. 

Do not operate without covers. Do not operate the instrument with covers or panels removed. 

Use Proper Fuse. Use only the specified type and rating fuse for this instrument. 

Avoid exposed circuit. Do not touch exposed junctions and components when the instrument is powered. 

Do not operate if in any doubt. If you suspect damage occurs to the instrument, have it inspected by qualified service personnel before further operations. 

Use your instrument in a well-ventilated area. Inadequate ventilation may cause an increasing of temperature or damages to the instrument. Please keep the instrument well ventilated, and inspect the air outlet and the fan regularly. 

Do not operate in wet conditions. To avoid short circuit inside the instrument or electric shock, never operate the instrument in a humid environment.
Do not operate in an explosive atmosphere.In order to avoid damages to the device or personal injuries, it is important to operate the device away from an explosive atmosphere.

Keep instrument surfaces clean and dry. To avoid the influence of dust or moisture in air, please keep the surface of device clean and dry.
Safety Terms and Symbols 

Safety Terms
Terms in this Manual. The following terms may appear in this manual: 
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  Warning: Warning indicates the conditions or practices that could result in injury or loss of life. 
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  Caution: Caution indicates the conditions or practices that could result in damage to this product or other property. 
Terms on the Product. The following terms may appear on this product: 
Danger: It indicates an injury or hazard may immediately happen. 

Warning: It indicates an injury or hazard may be accessible potentially. 

Caution: It indicates a potential damage to the instrument or other property might occur.
Safety Symbols

Symbols on the Product. The following symbol may appear on the product:
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	Hazardous Voltage
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	Refer to Manual
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	Protective Earth Terminal
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	Chassis Ground
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	Public Ground
	
	


Quick Start

Product Overview

The power meter adopts a dual-channel design. The rear panel (CH1) serves as the AC main input, AC supporting a maximum AC voltage and current of 600VAC/25kArms/75kApk. The front panel (CH2) supports both AC and DC sampling without harmonic analysis capability, optimized for DC sampling applications, and extends the DC input range up to 1000V voltage and 25kArms/75kApk current. With the dual-channel design, users can easily measure and calculate parameters such as voltage, current, power, frequency, crest factor, harmonics, integration, and efficiency.

Power Meter comes standard with RS232 communication interfaces, enabling remote control of the meter, and also features a USB host interface for saving measurement parameters directly to external storage devices, supporting long-term data recording. It offers 0.5% accuracy for voltage, current, and power measurements and includes extensive active power integration and energy accumulation functions. The power meter is widely used in testing household appliances, UPS systems, photovoltaic and wind power, charging stations, and energy storage applications.

Key Features:
Dual‑channel design (front & rear panels), supports external current sensors, suitable for high‑power applications:
Rear‑panel CH1 channel supports AC input: 3–600 VAC / 20 A RMS (built‑in) / 0–25 kA (external), with accuracy up to ±0.5%.

Front‑panel CH2 (AC/DC input) channel supports up to 0–1000 V / 0–20 A (built‑in) / 0–25 kA (external), with accuracy up to ±0.5%.

High-resolution 3.95-inch TFT color screen display

Simultaneous measurement of voltage, current, power, efficiency, and harmonics.

Integration and harmonic analysis up to the 63rd order.

Manual/auto range setting.

Customizable efficiency calculation support.

Multiple display modes: numeric, waveform, and bar graph.

USB support for screenshots and long-term data recording.

Standard digital communication interfaces RS232, supporting SCPI commands and Modbus protocol.

Panel Introduction
Front panel
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The components of the front panel are defined as follows:

	1
	Display Screen
	TFT color screen displays.

	2
	Menu select area
	Element: Channel switch key.
Auto: Automatic range setting.
VRange: Voltage range setting.
Irange: Current range setting.
Update Rate: Data update cycle.

	3
	Function key area 1
	Wiring: Wiring and efficiency settings.
Mode: Working mode settings.
PgUp: Page up to previous page.
ESC: Exit editing mode and return to the previous page.
Hold: Page hold key or pause.
PgDn: Page down to next page.

	4
	CH2 connection area
	CH2 Channel Voltage U+ U‑: Measured voltage of CH2. 

CH2 Channel Current I+ I‑: Measured current of CH2, with built‑in current sensor. 

CH2 Channel Current I+‑: Measured current of CH2, with external current sensor.

	5
	USB interface
	USB data interface.

	6
	Wave key
	Waveform display page switch key.

	7
	Numeric key
	Digital display page switch key.

	8
	Direction key, 
Enter key
	↑ ↓ ← →: Move the cursor up, down, left, and right.
Enter: Confirm parameter key.

	9
	Function key area 2
	Menu: Enter Menu.
Default: Return to the default value display page.
Record: Data recording with USB disk for a long time (effective when U disk is inserted).
Save: Save waveform screenshots to USB disk.
Local: Remote communication connection or short press Local to realize keyboard lock function; long press “Local” key to unlock the panel.
Shift: Reuse function.

	10
	Power button
	ON/OFF Switch.


Rear Panel
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	1
	CH1 voltage terminal
	The red V and black ± are the CH1 voltage sampling terminals.

	2
	External Sensor
	External BNC Current Sensor

	3
	Fuse
	Power supply fuse.

	4
	AC power terminal
	The power cord inputs AC power from this connection terminal to the input terminal.

	5
	Line switcher
	110V/220V AC voltage selection switch

	6
	CH1 current terminal
	The red I and black ± are the CH1 current terminals.

	7
	Ground 
	Protective grounding.

	8
	RS232 interface
	The computer can be connected via the RS232 port.


General Inspection

When you get a new digital power meter, it is recommended that you check the instrument according to the following steps.

1. Inspect the instrument for damage caused by shipping.

If the packaging carton or foam padding is severely damaged, please retain it until the instrument and accessories have passed electrical and mechanical tests.

2. Inspect the accessories.

A detailed list of supplied accessories is provided in "Appendix A: Accessories" of this manual. Refer to this information to check for missing accessories. If any accessories are missing or damaged, please contact your authorized distributor or local office.

3. Inspect the instrument.

If the instrument is damaged, malfunctions, or fails performance testing, please contact your authorized distributor or local office. If the instrument is damaged during shipping, please retain the packaging. Notify the shipping department and your authorized distributor. We will arrange for a repair or replacement.

Power-on inspection
Before operating the power meter, please ensure you understand the general safety precautions.

Before turning on the power, make sure the power supply voltage matches the supply voltage; otherwise, the instrument may be damaged. 

Please make sure that the main power plug is connected to a power socket with protective grounding. Do not use a power strip without protective grounding. Before operating the instrument, you should first ensure that the instrument is well grounded.

Please pay attention to the positive and negative pole markings and the maximum voltage and current limits before connecting the instrument, otherwise the instrument will be burned.
Press the power button on the front panel. The startup screen will appear.

Select Menu to enter the system settings interface. Press the SYSTEM INFO button to access the system information interface, which displays information such as the product model, serial number, and software version number.
Connection circuit 
The power meters can measure voltage, current, power, and other energy factors for various electrical products. This section describes typical circuit connection methods for practical applications of power meters.

Precautions before connection

When connecting a measurement circuit, always disconnect the power supply to the circuit to avoid the risk of electric shock during connection. 
Always connect the main power plug to a power outlet with a protective ground. Do not use a terminal block without a protective ground. Before connecting the circuit, ensure that the electronic load is properly grounded.

Never connect a current circuit to a voltage input terminal, or vice versa.

When insulating the measurement cable, ensure that the bare wires connected to the input terminal do not protrude from the terminal. Also, secure the screws of the input terminal to prevent the cable from becoming detached.

When connecting to the voltage input terminal, use a measurement cable with a safety rubber plug that does not expose bare wires. Ensure that the input terminal is securely fastened to prevent the cable from becoming detached.

Use a measurement cable that meets the rated voltage and current requirements and has a high voltage resistance and sufficient current capacity for the voltage and current being measured. When measuring a 20A current, use a copper wire with a cross-sectional area greater than 4mm2.

Connection Method

To achieve accurate measurements, consider the following when connecting the voltage and current input terminals. When measuring larger currents, connect the voltage input terminals closer to the load. When measuring smaller currents, connect the current input terminals closer to the load. The circuit diagram is shown below.
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Measuring larger currents           Measuring small currents
Voltage Input Terminals

The voltage terminals are safety rubber sockets. Insert the safety plug (without exposed wires) into the voltage input terminals.

Current Input Terminals

The current input terminals are binding posts. First, wrap the wire around the screw or insert the crimp terminal through the screw shaft, then hold the terminal knob and tighten the screw.

Measurement conditions

Setting the Crest Factor and Measurement Range

For the power meter, crest factor (CF) is defined as the ratio of the waveform peak value to the rated RMS value. CH1 has two ranges: CF = 3 and CF = 6. For example, when CF = 3 and the voltage range is 600V, the voltage peak that CH1 can measure is 600V * 3 = 1800V. The voltage range, current range, effective input range, and measurement accuracy depend on the crest factor setting.

Therefore, to perform accurate measurements, it is necessary to set the appropriate crest factor and voltage and current measurement ranges. The selected range is effective for different measurement modes. The voltage and current ranges of the power meter are shown in the table below.

	Voltage Range
	Current Range

	CH1:
CF=3:75V/150V/300V/600V/Auto

CF=6:37.5/75V/150V/300V/Auto

CH2:
50V/100V/500V/1000V/Auto
	CH1:
CF=3:0.2A/1A/4A/20A/Auto(Built-in)
CF=6:0.1A/0.5A/2A/10A/Auto(Built-in) 

CF=3:0.05V/0.1V/0.2V/0.5V/1V/2V/2.5V (Built-in)
CH2: 

0.2A/1A/4A/20A/Auto(Built-in)

0.05V/0.1V/0.2V/0.5V/1V/2V/2.5V(External)


Voltage and Current Ranges

The measurement range is set based on the levels of RMS value (CH1) and peak value (CH2). When inputting voltage or current signals into the input element, fixed range and auto range are available. The input of the external current sensor has a fixed range and only operates when CF=3.

Fixed Range
Select the desired range from multiple options. Once selected, the range will not change regardless of variations in the input signal level. 
For example, voltage range for CH1:

Crest factor(CF) of 3: The maximum option is "600V", and the minimum option is "75V".

Crest factor(CF) of 6: The maximum option is "300V", and the minimum option is "37.5V".

The ranges for other models will differ, please refer to the display of the specific model.

Auto-ranging Setting

Automatically switches range according to the magnitude of the input signal. The available ranges are the same as those for fixed ranges.
CH1 automatic range upshift principle: The range is upshifted when any of the following conditions are met.
Urms or Irms exceeds 110% of the currently set range.

CF= 3: The Upk or Ipk value of the input signal exceeds 330% of the currently set range.

CF= 6: The Upk or Ipk value of the input signal exceeds 660% of the currently set range.

CH1 automatic range downshift principle: The range is downshifted when all of the following conditions are met.

Urms or Irms is less than or equal to 70% of the lower range.
CF= 3: The Upk or Ipk value of the input signal is less than 300% of the lower range.

CF= 6: The Upk or Ipk value of the input signal is less than 600% of the lower range.

CH2 automatic range upshift principle: The range is upshifted when the following conditions are met.

Upk or Ipk exceeds 110% of the currently set range.

CH2 automatic range downshift principle: The range is downshifted when the following conditions are met.

Upk or Ipk is less than or equal to 70% of the lower range.

Note: When auto-ranging is selected, the range may change if the input waveform is a pulse waveform with an irregular period. In this case, select a fixed range.

Method 1: Operating Steps
Select Menu to enter the Range Setting interface.

In the CF Setting dropdown box, set the Crest Factor (CF). Two options are available: CF = 3 or CF = 6 (default is CF = 3).

After setting the crest factor, select the voltage range using the cursor directional keys to complete voltage range configuration.

After setting the voltage range, use the cursor directional keys to select the current sensor (internal or external). If the internal sensor is selected, further choose the current range to complete current‑range configuration. If an external sensor is selected, specify the range of the external sensor and set its scaling factor. For example, if the maximum output signal of the external current sensor is 2.2 V with a sensitivity of 0.22 V/A, select the (EXT) 2.5 V current range and set a custom sensor scaling factor of 0.22 V/A. Common current‑sensor scaling factors can be selected directly.
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Method 2: Operating Steps
On the main numeric display interface, if the user needs to set the voltage range, press the VRange key to open the voltage range dropdown box, or press the Auto key for quick automatic range setting. The Element key can be used to switch between different channels for voltage range configuration.

If the user needs to set the current range, press the IRange key to open the current range dropdown box, or press the Auto key for quick automatic range setting. The Element key can be used to switch between different channels for current range configuration. Display digits and power range of CH1 when CF=3.
When CF=3，Display digits and power range of CH1

	Voltage Range
	Current Range

	(V)
	0.200A
	1.000A
	4.000A
	20.000A

	75.00
	15.00W
	75.00W
	300.0W
	1500W

	150.0
	30.00W
	150.0W
	600.0W
	3000W

	300.0
	60.00W
	300.0W
	1200W
	6000W

	600.0
	120.0W
	600.0W
	2400W
	12000W


When CF=6，Display digits and power range of CH1

	Voltage Range
	Current Range

	(V)
	0.100A
	0.500A
	2.000A
	10.000A

	37.50
	3.75W
	18.75W
	70.00W
	375.0W

	75.00
	7.50W
	37.50W
	150.0W
	750.0W

	150.0
	15.00W
	75.00W
	300.0W
	1500W

	300.0
	30.00W
	150.0W
	600.0W
	3000W


Display digits and power range of CH2

	Voltage Range
	Current Range

	(V)
	0.200A
	1.000A
	4.000A
	20.000A

	50.00
	10.00W
	50.00W
	200.0W
	1000W

	100.0
	20.00W
	100.0W
	400.0W
	2000W

	500.0
	100.0W
	500.0W
	2000W
	10000W

	1000.0
	200.0W
	1000W
	4000W
	20000W


Setting the Data Update Rate

During measurement, the measurement interval determines the sampling data range. For the power meter, the measurement interval is defined by both the data update rate and the synchronization source. Each channel of the power meter uses its own voltage signal as the synchronization source. The synchronization source provides a reference signal for measurement, while the data update rate determines the update cycle of the sampled data. 

The power meter use the frequency measurement circuit to detect the period of the input signal, and the measurement interval is set as an integer multiple of the detected period. The power meter calculates the measurement values by averaging the sampled data within the measurement interval. However, for functions that measure the maximum values of voltage and current (Ipk+, Ipk-, Upk+, Upk-, Ucf, Icf), the measurement interval is based on the data update cycle. If the rising slope or falling slope appears only once or not at all within the data update cycle, the data update cycle itself is taken as the measurement interval. 

Note:
Data Update Rate: The time interval for each data update of the power meter. Available settings: 0.1 s, 0.2 s, 0.5 s, 1 s, 2 s, 5 s.

Data Update Cycle: The period of sampling data used for measurement functions, equal to the set data update rate.
Operating Steps

Method 1: Press the Update Rate key on the panel. The data update rate can be set cyclically among 0.1 s, 0.2 s, 0.5 s, 1 s, 2 s, and 5 s.

Method 2: Enter the System Settings via the Menu. In the Data UPDATE RATE dropdown box, select the desired update rate.
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Note: The data update rate is only effective in Normal Mode. When the input signal frequency is relatively low, the numeric display may become unstable and difficult to read. In this case, setting a larger data update rate allows averaging over a longer period to obtain more stable measurement data.

Valid Input Range and Abnormal Data Handling

For the power meters:

Channel 1 and Channel 2 rated voltages are 600 Vrms and 1000 Vpk, supporting continuous input of 700 Vrms and 1100 Vpk, with maximum rated peak voltages of 1800 V (at CF=3) and 1100 V.

The rated maximum of the built-in current sensor is 20 Arms and 20 Apk, supporting continuous input of 24 Arms and 24 A, with maximum rated peak currents of 60 A (at CF=3) and 30 A.

The rated maximum of the external current sensor is 25 kA rms, with a peak current up to 75 kA pk.
Abnormal Data Handling:

When voltage Urms or current Irms is zero, related parameters including active power P, apparent power S, reactive power Q, power factor λ, phase angle φ, voltage frequency fU, voltage peak Upk, current peak Ipk, voltage crest factor CfU, and current crest factor CfI are not calculated, and “----” is displayed.

When the voltage or current exceeds the measurement range (CH1 voltage Urms exceeds 1.2 times the current range, and current Irms exceeds 1.2 times the current range; CH2 voltage Upk exceeds 1.1 times the current range, and current Ipk exceeds 1.2 times the current range), “OVER” is displayed, indicating overrange. In this case, P, S, and Q also display “OVER”.

For DC input, if no frequency is detected, the panel displays “----”. For AC input, if fU frequency is outside the 38–72 Hz range, the panel displays “ERR”.

When |U1pk+| or |U1pk-| exceeds CF × current voltage range, the panel displays “OVER”.

When |I1pk+| or |I1pk-| exceeds CF × current current range, the panel displays “OVER”.

When |U2pk+| or |U2pk-| exceeds 1.1 times the current voltage range, the panel displays “OVER”.

When |I2pk+| or |I2pk-| exceeds 1.2 times the current current range, the panel displays “OVER”.

When active power values used for efficiency calculations (η1, η2) overflow, the panel displays “OVER”.

The range of THD (Total Harmonic Distortion) is [0, 400%]. If the value exceeds this range, the panel displays “ERR”.

General Measurement Functions

General Information 
The power meters provide a comprehensive set of basic electrical measurement functions, accurately measuring voltage, current, power, power factor, harmonics, frequency, and distortion factor. The reliable data serves as a scientific basis for analyzing the energy quality and performance of equipment.

The variable naming convention of the digital power meter is: parameter type + channel number + description. For example, U1rms and U2rms represent the RMS voltage of Channel 1 and Channel 2, respectively.

The phase angle represents the phase shift between channel voltage and current. φ1 denotes the phase angle between U1 and I1. A negative value means current leads voltage, while a positive value means current lags voltage.
General Measurement Parameters

	Symbol
	Description
	Symbol
	Description
	Symbol
	Description

	P
	Active Power [W]
	Irms
	RMS Current [A]
	Urms
	RMS Voltage [V]

	Q
	Reactive Power [var]
	S
	Apparent Power [VA]
	λ
	Power Factor

	Upk+
	Positive Voltage Peak [V]
	Ipk+
	Positive Current Peak [A]
	CfU
	Voltage Crest Factor

	Upk-
	Negative Voltage Peak [V]
	Ipk-
	Negative Current Peak [A]
	CfI
	Current Crest Facto

	ф
	Phase Angle between Voltage and Current
	fU
	Voltage Frequency [Hz]
	η
	Efficiency


Formulas for General Measurement Parameters：
	Measurement function[unit]
	Formula

	Urms [V]
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	Upk+ [V]
	The maximum u(n) for every data update u(n)

	Upk- [V]
	The minimum u(n) for every data update u(n)

	Ipk+ [V]
	The maximum u(n) for every data update i(n)

	Ipk- [V]
	The minimum u(n) for every data update i(n)

	CfU
	CfU=MAX(|Upk+|,|Upk-|)/Urms

	CfI
	CfI=MAX(|Ipk+|,|Ipk-|)/Irms


Wiring Methods
The power meter adopts 1P2W wiring method. See the figure below for detailed wiring.
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Connect CH1 voltage to the rear panel U ± terminals and CH1 current to the rear panel I ± terminals. Connect CH2 voltage to the front panel U+ U- terminals and CH2 current to the front panel I+ I- terminals. Alternatively, connect external current sensors to the front and rear panels via BNC cables or BNC-to-banana adapters.

Operating Steps:
The user can press the Wiring button to enter the wiring mode and efficiency setting interface.

Press the PgDn key on the power meter to enter the efficiency formula setting interface. Select the option box for the η1 formula to choose the desired calculation formula flexibly, such as η1=P1/P2, η1=P2/P1, η1=P2/Udef1, etc.
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Setting Measurement Functions
The power meter features 3 numeric display styles and 1 waveform display style for regular measurements, with up to 9 pages for each style. To highlight key measurement parameters, switch freely to 4ITEM or 8ITEM display mode. To view all parameters on a single screen, select 16ITEM mode. Switch to waveform display mode to observe channel waveforms.

The default parameters displayed in the 4-ITEM, 8-ITEM, and 16-ITEM display modes are shown in the table below.

	4ITEM

	1
	2
	3
	4
	5
	6
	7

	U1rms
	Q1
	U1pk+
	CfU1
	U2rms
	U2rms
	U2pk+

	I1rms
	λ1
	U1pk-
	CfI1
	I2rms
	λ2
	U2pk-

	P1
	φ1
	I1pk+
	η1
	P2
	--
	I2pk+

	S1
	fU1
	I1pk-
	--
	S2
	fU2
	I2pk-


	8ITEM

	1
	2
	3
	4
	5
	6

	U1rms
	U1pk+
	U2rms
	---
	---
	---

	I1rms
	U1pk-
	I2rms
	---
	---
	---

	P1
	I1pk+
	P2
	---
	---
	---

	S1
	I1pk-
	S2
	---
	---
	---

	Q1
	CfU1
	---
	---
	---
	---

	λ1
	CfI1
	λ2
	---
	---
	---

	φ1
	η1
	---
	---
	---
	---

	fU1
	---
	fU2
	---
	---
	---

	U1rms
	U1pk+
	U2rms
	---
	---
	---

	I1rms
	U1pk-
	I2rms
	---
	---
	---

	P1
	I1pk+
	P2
	---
	---
	---

	S1
	I1pk-
	S2
	---
	---
	---

	Q1
	CfU1
	---
	---
	---
	---

	λ1
	CfI1
	λ2
	---
	---
	---

	φ1
	η1
	---
	---
	---
	---

	fU1
	---
	fU2
	---
	---
	---


	16ITEM

	1
	2
	3
	4
	5
	6

	U1rms
	U2rms
	---
	---
	---
	---

	I1rms
	I2rms
	---
	---
	---
	---

	P1
	P2
	---
	---
	---
	---

	S1
	S2
	---
	---
	---
	---

	Q1
	---
	---
	---
	---
	---

	λ1
	λ2
	---
	---
	---
	---

	φ1
	η1
	---
	---
	---
	---

	---
	---
	---
	---
	---
	---

	fU1
	
	---
	---
	---
	---

	U1pk+
	U2pk+
	---
	---
	---
	---

	U1pk-
	U2pk-
	---
	---
	---
	---

	I1pk+
	I2pk+
	---
	---
	---
	---

	I1pk-
	I2pk-
	---
	---
	---
	---

	CfU1
	---
	---
	---
	---
	---

	CfI1
	---
	---
	---
	---
	---

	---
	---
	---
	---
	---
	---


Operating Steps
Press the Mode button or access it from the Menu to enter the mode settings, enabling the power meter to operate in normal measurement mode. By pressing the Mode button continuously, the power meter can cycle through Normal Mode, Harmonic Mode, and Integration Mode.

Press the Wiring button or access it from the Menu to enter the wiring mode and efficiency formula settings. Select the appropriate wiring configuration and set the efficiency formula.
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Press the Esc key to return to the main numerical display page of normal measurement. The user can press the Numeric button to switch between numerical display interfaces. By pressing this button repeatedly, the display will cycle through different styles: 4-ITEM, 8-ITEM, and 16-ITEM.
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To modify display items, the user can move the cursor using the Up/Down arrow keys to the target position (e.g., the second row to display P2), then press Enter to enter edit mode. At this point, P2 can be directly selected from the dropdown menu and confirmed with Enter. Alternatively, the user may select the same parameter type from another channel (e.g., P1), then press the Element key to switch the channel to P2. 

Repeat the above step to complete the configuration of all parameters to be displayed.

If additional display variables are needed, press the PgUp / PgDn keys to switch to different display pages. The power meters support up to nine pages. Parameters on different pages can be viewed by toggling with the PgUp / PgDn keys.

To view voltage and current waveforms, press the Wave button to switch to the waveform display interface. From the dropdown menu, can select the channel to enable waveform display. 
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Press the Numeric key again to return to the numeric display interface.
Alarm Setting Function 
In certain automated testing scenarios, the power meters provide upper and lower limit alarm functions for voltage, current, and power parameters. Users can configure these alarm limits through the Application Settings menu.
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Users can enable or disable the alarm function by switching the Alarm On/Off setting. To reduce false alarms caused by sampling noise, the alarm delay count can be set. An alarm will only be triggered if the monitored parameter exceeds the set limit continuously for the number of times specified by the alarm delay count.

For example, if CH1 is set with an alarm delay count of 5, and the voltage of CH1 exceeds the upper limit for 5 consecutive data update cycles, a pop-up will indicate an upper voltage limit alarm, and the alarm indicator in the status bar will turn red. Pop-up alarm messages must be manually cleared by the user. Status bar alarm indicators update in real time according to the current measurement. If the sampled value returns to the normal range, the status bar will revert to black.

For the 1P2W single-phase wiring mode or built-in current sensor, the CH1 voltage, current, and power alarm information is determined based on |U1Arms|, |I1Arms|, and |P1A|. In this case, the maximum alarm thresholds for voltage, current, and power are 660 V, 22 A, and 13.2 kW, with the minimum value being 0.

For CH2, the maximum alarm thresholds for voltage |U2rms|, current |I2Arms|, and power |P2| are 1100 V, 22 A, and 22 kW, respectively, with the minimum value being 0.

When both the upper and lower limits are set to zero, it indicates that the parameter is not subject to alarm monitoring.

Operating Steps
Press the Menu key to enter the Alarm Settings interface.

The user can independently configure the alarm information for CH1 and CH2. When a specific parameter needs to be monitored, its corresponding upper and lower alarm limits must be set to a non-zero value by using the Up/Down/Left/Right arrow keys, along with the Enter and Esc keys.

To monitor a specific parameter, use the cursor keys (Up/Down/Left/Right), Enter, and Esc to set the corresponding upper and lower alarm limits to non-zero values.

After setting the limits, turn on the Alarm Switch to start monitoring voltage, current, and power limits.
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Screenshot and Data Logging Functions
The power meters provide screenshot saving and USB flash drive data logging functions, enabling users to conveniently store test data. The key information for USB data logging is as follows:

USB flash drive requirements: FAT32 format, capacity ≤ 32 GB

Status indicators: Gray (not connected) → White (recognized) → Flashing (read/write in progress)

Logging steps: Insert USB drive → System Settings → USB Logging → Set Length → Start Logging
Logging status: Status bar displays "T00:00:00" + USB flashing

Termination conditions: Manual stop / Reaching set duration / Automatic stop after 72 hours / Insufficient space / USB removed

File storage: Named as PA_DATA_<time>.csv. When data exceeds 1 million records, a new file is automatically created.

Data format: Includes timestamp + parameters of the current page

Screenshot Saving Steps:
Insert the USB flash drive and ensure it is successfully recognized.

During operation, press the Hold button to pause screen updates.

Press and hold the Save button to capture and save a screenshot.

Data Logging Steps:

Insert the USB flash drive and ensure it is successfully recognized.

Press the Menu button to enter the USB Logging Settings interface.

Use the edit function to select the parameters to be recorded and set the recording duration. After saving the settings, exit to the main screen.

Press the Record button to start logging. Press the Record button again to stop logging.
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Harmonics Measurement Function
This chapter describes the harmonics measurement features and operation of the power meters.
Basic Concepts 
The CH1 channel of the digital power meter supports harmonics analysis. Its full-bandwidth passband is 2.8 kHz, the measurable fundamental frequency range is 45–66 Hz, and the maximum number of measurable harmonics is 63. Harmonics measurement error is 1%.

The meter can measure active power, reactive power, power factor, and total harmonic distortion (THD) for voltage, current, and each harmonic order. Harmonic parameters can be displayed in lists or bar charts, allowing intuitive analysis of the measurement results.

Harmonics Measurement Operation 
The harmonics measurement mode provides three main display interfaces: list, numeric, and bar chart.

Use list mode to view multiple harmonic orders simultaneously.

Use numeric mode to focus on specific harmonic values of voltage, current, or power.

Use bar chart mode for a visual representation of harmonic orders.
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Operating Steps
Press the Mode key to set the power meter to Harmonics Mode.

Press the Numeric key to switch between list and numeric displays. Repeatedly pressing this key cycles between the two display styles.

List Display: (diagram or table of multiple harmonic orders)
Single Harmonic Numeric Display: (diagram showing single harmonic value)
Use the Up/Down arrow keys to move the cursor to the desired setting position. For example, to select the channel for CH1, press Enter to enter edit mode. At this point, other channel options can be selected from the dropdown menu and confirmed by pressing Enter. Alternatively, to set the number of harmonics to read, press Enter to enter edit mode, then use the Up/Down keys to set the harmonic count from 1 to 63. Press Enter to confirm and exit the harmonic count setting.

Press the Numeric button to switch to the single-harmonic numerical display interface. When the cursor is on the harmonic count setting, press Enter to confirm and set the single-harmonic count from 1 to 63. After setting, press Enter to confirm and exit the single-harmonic count setting.

If the user wants to view the voltage or current harmonic waveforms, press the Wave button to switch to the harmonic mode display interface. By selecting the parameters, the user can switch between voltage and current to view individual harmonic displays.
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To return to the numeric display interface, press the Numeric key again to switch back to the previous numeric display screen.
Integration Measurement Function
This chapter describes the integration measurement features and operation methods of the power meters. 

Basic Concepts 
Both CH1 and CH2 channels of the power meter can perform:

Active power integration (Wh)
Current integration (Ah)
Integration supports automatic or fixed range settings. During integration, the display can show Wh (WP), Ah (q), and integration time (Time), as well as measurement and calculated values from normal or harmonics measurement modes.

Integration formulas:
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Current integration uses RMS calculation mode.

Maximum/minimum display values:

Active power (WP): ±999,999 MWh

Current (q): ±999,999 MAh

Time: 10,000 hours
The maximum display resolution is 999,999. When the integration value reaches 1,000,000 counts, the decimal point automatically moves. For example, 999.999 mWh + 0.001 mWh = 1.00000 Wh.

Integration Overflow Display
Integration stops and retains the current integration value and time when:

Integration time reaches the maximum (10,000 hours)

WP or q reaches the maximum/minimum display value

Integration with Measurement Limits
If instantaneous voltage or current exceeds the AD circuit range, these values are treated as the upper/lower limit of the range.

Integration Modes
The power meter supports Manual Integration Mode and Standard Integration Mode.

Manual Integration Mode

Integration continues from start to stop. However, integration stops automatically if:

Maximum integration time (10,000 hours) is reached

Integration value reaches maximum/minimum display value
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Standard Integration Mode

Integration time is set via a relative timer. Integration stops when:

The set time expires

Maximum/minimum display value is reached before the timer ends
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Integration Operation 
Start, Stop, Hold, Reset

Users can start, stop, reset, or hold integration via front panel keys or communication commands to control the CH1 and CH2 together.
Start: Shift + Wiring or panel Start key

Stop: Shift + Hold or panel Stop key

Reset: Shift + Esc or panel Reset key

Hold: Hold key (locks the display while integration continues)

Integration Main Interface

[image: image40.png]Al

PARAMETER INTEGRATION MODE
CH1 CH1 CH2

SeTTIME st | staT

0000: 01: 00

CH2 WP1 wP2

ez 0.00000 mWh o 00000 mWh
T«

SETTIME 0.00000 mAh 0 00000 mAh

00000100 | TESTTIME: 0000:00:00 TEST TIME: 0000:00:00




The main interface includes:

Parameter Setting Area:

Mode: Manual or Standard integration

Set Time: max 9999:59:59 (10,000 hours)

Display Control Area:

Status: Run, Stop, Hold, Reset

Test Time: 9999:59:59 (current integration duration)

Wh Integration: max 999,999 MWh

Ah Integration: max 999,999 MAh

Basic Operations

Set voltage/current range via Mode key or Menu to enter integration mode.

Enter integration mode via Mode key or Menu.

During integration, normal and harmonic measurements can be displayed. Status bar shows integration running symbol.

Hold Function: Press Hold to freeze the display while integration continues.

Run Function: Press Shift+Wiring or panel Start key to begin integration.
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Stop Function: Press Shift+Hold or panel Stop key to pause integration.
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Reset Function: Press Shift+Esc or panel Reset key to clear integration values.
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Manual Integration Operation

Set Mode via key or Menu to enter integration mode.

Select Manual from the mode dropdown (time setting disabled).

Press Shift+Wiring or panel Start key to begin integration.

Stop integration via Shift+Hold or panel Stop key.

Standard Integration Operation

Set Mode via key or Menu to enter integration mode.

Select Normal mode from the dropdown.

Set the timer (max 9999:59:59).

[image: image44.jpg]PARAMETER
CcHL
MODE

SETTIME
0000:00: 30
CH2

[

SETTIME
0000: 01: 00

Al
INTEGRATION MODE
CH1 CH2
REsET. L
smarT staT

wP1 wp2
350.222 mWh | 0.00000 mWh
al a2

9.49546 mAh | 0.00000 mAh

TEST TIME: 0000:00:30| TEST TIME: 0000:00:00




Start integration via Shift+Wiring or panel Start key.

Stop integration manually if needed via Shift+Hold or panel Stop key.
Troubleshooting

If the instrument remains black with no display after pressing the power switch, please proceed as follows:

Check whether the power connector is properly connected.

Check whether the fuse below the power input socket is correctly selected and intact (use a flat-head screwdriver to open if necessary).

After completing the above checks, restart the instrument.

If the problem persists, please contact our company for service assistance.

Technical Specifications
The instrument must be operated continuously for more than 30 minutes at the specified operating temperature to meet the following specifications.
	Model
	OWH9822

	Display
	TFT 3.95 inch Screen

	Update rate
	0.1s、0.2s、0.5s、1s、2s、5s

	Measurement items
	V, A, W, VA, var, λ, fU, CfU, CfI, Upk, Ipk, THD

	Measurement method
	CH1

1ph-AC
	CH2

1ph-AC+DC

	Voltage measurement range
	3.0V-600Vrms   
	0V-1000Vpk*1

	Voltage range
	CF=3：75V/150V/300V/600V

CF=6：37.5/75V/150V/300V
	50V/100V/500V/1000V

	Voltage accuracy
	±( 0.4% of reading + 0.1% of range + 1 digit)

	Voltage resolution
	0.01V/0.1V

	Current measurement range
	0~20.0Arms(built-in)/

0-25kA(external)
	0-20Apk*1(built-in)/

0-25kA(external)

	Current range
	CF=3:0.2A/1A/4A/20A(built-in)

CF=6:0.1A/0.5A/2A/10A(built-in)

CF=3:0.05V/0.1V/0.2V/0.5V/1V/

2V/2.5V(external)
	0.2A/1A/4A/20A(built-in)
0.05V/0.1V/0.2V/0.5V/1V/2V

/2.5V(external)

	Current accuracy
	±( 0.4% of reading + 0.1% of range + 1 digit)

	Current resolution
	0.001A

	Range Switching
	Auto/Manual

	Power Range
	0~36kW
	0~20kW

	Power accuracy
	±( 0.4% of reading +0.1% of range +1 digit)

	Power Resolution
	0.01W/0.1W/1W

	Power Factor Range
	-1.000~1.000

	Power Factor Accuracy
	±( 0.4% of reading +0.1%* of reading +1 digit)
	----

	Operating Frequency
	40Hz~70Hz
	DC+40Hz-70Hz

	Frequency Accuracy
	±(0.1% of reading + 1 digit)
	----

	Harmonic Analysis
	2.8 kHz bandwidth, up to 63rd order
	----

	Integration Function
	Watt-hour (Wh), Ampere-hour (Ah)

	Warm-up Time
	>30 min

	Peak Current
	60A
	30A

	Maximum Continuous Input
	Vrms=700V; Irms=24A
	Vpk=1100V; Ipk=24A

	Interface
	RS232

	Communication Protocol
	Standard SCPI ， Modbus protocol

	Baud Rate (bps)
	4800、9600(default)、19200、38400、57600，115200,

	Display Hold
	YES

	Mute
	YES

	Key Lock
	YES

	Power Supply
	AC 100V~240V, 50/60Hz

	Operation environment
	18℃~28℃，30%~75%RH (When operating temp. is outside 18 ℃ ~ 28 ℃, add temp. coefficient: 0.05% of reading / ℃）

	Storage Environment
	-10℃~50℃，80% RH (non-condensing)

	Operating Altitude
	≤2000 meters

	Weight
	1.3kg

	Dimensions
	235mm(D) × 110mm(H) × 295mm(W)


*1: DC peak maximum value.

Unless otherwise specified, all above parameters apply to the built-in sensors. The measurement ranges stated herein are valid when voltage and current are above 20% of their rated ranges. Refer to Section 4.1 for the definition of 1-digit error.
Adjustment Interval: A calibration interval of one year is recommended.

Appendix
Appendix A: Accessories
 (Images for reference only, subject to the actual product.)
	Standard Accessories:
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Power cord
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User manual
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Banana plug to crocodile clip test leads
	
	

	Optional Accessories:
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Banana head
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RS232 Cable
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Terminal Box Accessories
	
	


Appendix B: Maintenance and Cleaning

General Maintenance
Do not store or place the instrument in locations where the LCD display is exposed to direct sunlight for extended periods.
Caution: Do not allow sprays, liquids, or solvents to come into contact with the instrument to avoid damage.

Cleaning
Inspect the instrument regularly depending on usage. Clean the exterior of the instrument following the steps below:
Wipe off surface dust with a soft cloth. When cleaning the LCD display, take care not to scratch the transparent LCD protective cover.

Wipe the instrument with a damp but not dripping soft cloth. Ensure the power is disconnected before cleaning. Mild detergent or clean water may be used. Do not use any abrasive chemical cleaners, which may damage the instrument
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	Warning: 
Before reconnecting power and operating the instrument, make sure it is completely dry to prevent electrical short circuits or personal injury caused by residual moisture.
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