SP Series Single Output DC Power Supply Technical Specifications

The instrument must be operated continuously for more than 30 minutes at the specified operating temperature to achieve the following specifications:

|  |  | Parameter |  |  |  |  |  |  |  | Description |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | SP3051 | SP6053 | SP3101 | SP3103 | SP6101 | SP6103 | SP8206 | SP8207 |  |
| Input Parameter | Work Voltage Range | 100-120Vac | 100-120Vac | 100-120Vac | 100-120Vac | 100-120Vac | 100-120Vac | 100-240Vac | 200-240Vac | 110V± 10% |
|  |  | 200-240Vac | 200-240Vac | 200-240Vac | 200-240Vac | 200-240Vac | 200-240Vac |  |  | 220V± 10% |
|  | Work Frequency Range | 50-60Hz | 50-60Hz | 50-60Hz | 50-60Hz | 50-60Hz | 50-60Hz | 50-60Hz | 50-60Hz |  |
|  | Efficiency | 80% | 80% | 80% | 80% | 80% | 80% | ≥83% | -- | 110Vac，full load |
|  |  | 85% | 85% | 85% | 85% | 85% | 85% | ≥88% | ≥88% | 220Vac，full load |
|  | Full load input current | ≤2.8A | ≤5.6A | ≤2.8A | ≤5.6A | ≤2.8A | ≤5.6A | ≤6.6A | -- | 110Vac |
|  |  | ≤1.4A | ≤2.8A | ≤1.4A | ≤2.8A | ≤1.4A | ≤2.8A | ≤3.2A | ≤3.8A | 220Vac |
|  | No-load input current | ≤200mA | ≤300mA | ≤300mA | ≤300mA | ≤300mA | ≤300mA | ≤200mA | -- | 110Vac |
|  |  | ≤100mA | ≤150mA | ≤150mA | ≤150mA | ≤150mA | ≤150mA | ≤100mA | ≤150mA | 220Vac |
| Rated Output (0℃-40℃) | Voltage | 0-30V | 0-60V | 0-30V | 0-30V | 0-60V | 0-60V | 0-80V | 0-80V |  |
|  | Overvoltage Protection | 31V | 61V | 31V | 31V | 61V | 61V | 85V | 85V |  |
|  | Current | 0-5A | 0-5A | 0-10A | 0-10A | 0-10A | 0-10A | 0-20A | 0-20A |  |
|  | Power | 150W | 300W | 150W | 300W | 150W | 300W | 600W | 700W |  |
| Load Regulation (%of output+offset) | Voltage | ≤30mV | ≤30mV | ≤30mV | ≤30mV | ≤30mV | ≤30mV | ≤30mV | ≤30mV | 0-100% |
|  | Current | ≤30mA | ≤30mA | ≤30mA | ≤30mA | ≤30mA | ≤30mA | ≤30mA | ≤30mA | 90%-10% Rated Voltage |
| Line Regulation (%of output+offset) | Voltage | ≤20mV | ≤20mV | ≤20mV | ≤20mV | ≤20mV | ≤20mV | ≤20mV | ≤20mV |  |
|  | Current | ≤20mA | ≤20mA | ≤20mA | ≤20mA | ≤20mA | ≤20mA | ≤20mA | ≤20mA |  |
| Setting Resolution | Voltage | 10mV | 10mV | 10mV | 10mV | 10mV | 10mV | 10mV | 10mV |  |
|  | Current | 1mA | 1mA | 1mA | 1mA | 1mA | 1mA | 1mA | 1mA |  |
| Readback Resolution | Voltage | 10mV | 10mV | 10mV | 10mV | 10mV | 10mV | 10mV | 10mV |  |
|  | Current | 1mA | 1mA | 1mA | 1mA | 1mA | 1mA | 1mA | 1mA |  |
| Setting Accuracy (within 12 months) (25℃±5℃) (%of output+offset) | Voltage | ≤0.1%± 20mV | ≤0.1%± 20mV | ≤0.1%± 20mV | ≤0.1%± 20mV | ≤0.1%± 30mV | ≤0.1%± 30mV | ≤0.1%± 20mV | ≤0.1%± 20mV |  |
|  | Current | ≤0.1%± 10mA | ≤0.1%± 10mA | ≤0.1%± 10mA | ≤0.1%± 10mA | ≤0.1%± 10mA | ≤0.1%± 10mA | ≤0.1%± 20mA | ≤0.1%± 20mA |  |
| Readback Accuracy (25℃±5℃) (%of output+offset) | Voltage | ≤0.1%± 20mV | ≤0.1%± 20mV | ≤0.1%± 20mV | ≤0.1%± 20mV | ≤0.1%± 30mV | ≤0.1%± 30mV | ≤0.1%± 20mV | ≤0.1%± 20mV |  |
|  | Current | ≤0.1%± 10mA | ≤0.1%± 10mA | ≤0.1%± 10mA | ≤0.1%± 10mA | ≤0.1%± 10mA | ≤0.1%± 10mA | ≤0.1%± 20mA | ≤0.1%± 20mA |  |
| Ripple and noise (20Hz-20MHz) | Voltage(Vp-p) | ≤30mVp-p | ≤30mVp-p | ≤30mVp-p | ≤30mVp-p | ≤50mVp-p | ≤50mVp-p | ≤50mVp-p | ≤50mVp-p |  |
|  | Current(rms) | ≤3mVrms | ≤3mVrms | ≤3mVrms | ≤3mVrms | ≤5mVrms | ≤5mVrms | ≤20mVrms | ≤20mVrms |  |
|  | Current | ≤30mAp-p | ≤30mAp-p | ≤30mAp-p | ≤30mAp-p | ≤30mAp-p | ≤30mAp-p | ≤20mArms | ≤20mArms |  |
| Output temperature coefficient (0℃-40℃) (%of output+offset) | Voltage | ≤0.3%± 20mV | ≤0.3%± 10mV | ≤0.3%± 10mV | ≤0.3%± 10mV | ≤0.3%± 20mV | ≤0.3%± 20mV | 100ppm/℃ |  |  |
|  | Current | ≤0.3%± 10mA | ≤0.3%± 20mA | ≤0.3%± 20mA | ≤0.3%± 20mA | ≤0.3%± 20mA | ≤0.3%± 20mA | 200ppm/℃ |  |  |
| Readback temperature coefficient (%of output+offset) | Voltage | ≤0.3%± 20mV | ≤0.3%± 10mV | ≤0.3%± 10mV | ≤0.3%± 10mV | ≤0.3%± 20mV | ≤0.3%± 20mV | 100ppm/℃ |  |  |
|  | Current | ≤0.3%± 10mA | ≤0.3%± 20mA | ≤0.3%± 20mA | ≤0.3%± 20mA | ≤0.3%± 20mA | ≤0.3%± 20mA | 200ppm/℃ |  |  |
| Transient response time | Time | ≤1.0ms | ≤1.0ms | ≤1.0ms | ≤1.0ms | ≤1.0ms | ≤1.0ms | ≤5.0ms | ≤5.0ms | (50%-100% Rated Load) (SP8206/SP8207 model: 10%-90% Rated Load) |

General Specifications

Display

| Characteristics | Description |
| --- | --- |
| Display Type | 3.7 inches color LCD |
| Resolution | 480 × 320 pixels |
| Color | 65536 color, TFT |

Power Supply

| Characteristics | Description |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Power Supply | 110 VAC ± 10% or 220 VAC ± 10% power supply according to the label on the left side of the power input socket on the rear panel of the instrument); AC input 50/60 Hz |  |  |  |  |  |  |  |  |
| Fuse | Parameter | SP3051 | SP6053 | SP3101 | SP3103 | SP6101 | SP6103 | SP8206 | SP8207 |
|  | 110 V | 250 V, F5A | 250 V, F10A | 250 V, F5A | 250 V, F10A | 250 V, F5A | 250 V, F10A | 250 V, F10A | -- |
|  | 220 V | 250 V, F3A | 250 V, F5A | 250 V, F3A | 250 V, F5A | 250 V, F3A | 250 V, F5A | 250 V, F10A | 250 V, F10A |

Environment

| Characteristics | Description |
| --- | --- |
| Temperature | Working temperature: 0℃～40℃ Storage temperature: -20℃～60℃ |
| Relative Humidity | ≤90% |
| Height | Operation: 3,000 meters Non-operation: 15,000 meters |
| Cooling | Fan cooling |

Mechanical Specifications

| Characteristics | Description |  |
| --- | --- | --- |
| Dimension | 290mm(D)× 117mm(W)× 194mm(H) |  |
| Weight | SP3051/SP6053/SP3101/SP3103/SP6101/SP6103 | Approx. 3.0kg |
|  | SP8206/SP8207 | Approx. 2.0kg |

Safety Specifications

| Characteristics | Description |
| --- | --- |
| Withstand Voltage (Output to Ground) | 500VDC |

V1.1.5