SPE50 Series Single Output Power Supply Technical Specifications
The instrument must be operated continuously for more than 30 minutes at the specified operating temperature to achieve the following specifications.
	Model
	SPE50024
	SPE50025

	Input Characteristics                                                                                        

	Supply Voltage
	100V-240V

	Input Frequency
	45~65Hz

	Input Fuse
	250V,F10A

	Rated Output	                                                                                            

	Voltage
	0~500V
	0~500V

	Current
	0~2A
	0~2A

	Power
	400W
	500W

	Regulation (CV)					

	Load
	≤200mV

	Line
	≤200mV

	Regulation (CC)					

	Load
	≤10mA

	Line
	≤10mA

	Ripple & Noise (Noise bandwidth 20MHz, ripple bandwidth 1MHz, connect 10uF electrolytic capacitor in parallel with 0.1uF ceramic capacitor to the output terminal for testing)

	Voltage (Vp-p)
	≤500mVp-p
	≤500mVp-p

	Current (rms)
	≤20mA
	≤20mA

	Setting Resolution

	Voltage
	10mV
	10mV

	Current
	1mA
	1mA

	Readback Resolution

	Voltage
	10mV
	10mV

	Current
	1mA
	1mA

	Setting Accuracy

	Voltage
	≤0.05%±200mV

	Current
	≤0.05%±20mA
	≤0.05%±20mA

	Readback Accuracy

	Voltage
	≤0.05%±200mV

	Current
	≤0.05%±20mA
	≤0.05%±20mA

	Response time

	Transient recovery time (10%~90% rated load)
	≤5.0ms

	Protection function

	OVP
	0~510V
	0~510V

	OCP
	0~3A
	0~3A

	OTP
	85℃
	85℃

	Temperature coefficient of output

	Voltage
	100ppm/℃

	Current
	200ppm/℃

	Temperature coefficient of readback value

	Voltage
	100ppm/℃

	Current
	200ppm/℃

	Display

	Display Type
	2.8-inch Color LCD Display (IPS Panel)

	Resolution
	240 × 320 pixels

	Color
	65536 color,TFT

	Environment

	Temperature
	Working Temperature:0℃～40℃
Storage Temperature:-20℃～60℃

	Relative Humidity
	≤90%RH; no condensation

	Height
	2,000 meters

	Cooling
	Fan cooling, temperature intelligent speed control

	Others

	Communication Port
	[bookmark: _GoBack]USB, LAN(optional), DB9:RS232(optional), compatible with the SCPI communication protocol

	Dimension
	82mm(W)× 142mm(H)×226mm(D)

	Weight
	Approx. 1.6kg
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