SPE50 Series Single Output Power Supply Technical Specifications

The instrument must be operated continuously for more than 30 minutes at the specified operating temperature to achieve the following specifications.

| Model | SPE50024 | SPE50025 |
| --- | --- | --- |
| Input Characteristics |  |  |
| Supply Voltage | 100V-240V |  |
| Input Frequency | 45~65Hz |  |
| Input Fuse | 250V,F10A |  |
| Rated Output |  |  |
| Voltage | 0~500V | 0~500V |
| Current | 0~2A | 0~2A |
| Power | 400W | 500W |
| Regulation (CV) |  |  |
| Load | ≤200mV |  |
| Line | ≤200mV |  |
| Regulation (CC) |  |  |
| Load | ≤10mA |  |
| Line | ≤10mA |  |
| Ripple & Noise (Noise bandwidth 20MHz, ripple bandwidth 1MHz, connect 10uF electrolytic capacitor in parallel with 0.1uF ceramic capacitor to the output terminal for testing) |  |  |
| Voltage (Vp-p) | ≤500mVp-p | ≤500mVp-p |
| Current (rms) | ≤20mA | ≤20mA |
| Setting Resolution |  |  |
| Voltage | 10mV | 10mV |
| Current | 1mA | 1mA |
| Readback Resolution |  |  |
| Voltage | 10mV | 10mV |
| Current | 1mA | 1mA |
| Setting Accuracy |  |  |
| Voltage | ≤0.05%±200mV |  |
| Current | ≤0.05%±20mA | ≤0.05%±20mA |
| Readback Accuracy |  |  |
| Voltage | ≤0.05%±200mV |  |
| Current | ≤0.05%±20mA | ≤0.05%±20mA |
| Response time |  |  |
| Transient recovery time (10%~90% rated load) | ≤5.0ms |  |
| Protection function |  |  |
| OVP | 0~510V | 0~510V |
| OCP | 0~3A | 0~3A |
| OTP | 85℃ | 85℃ |
| Temperature coefficient of output |  |  |
| Voltage | 100ppm/℃ |  |
| Current | 200ppm/℃ |  |
| Temperature coefficient of readback value |  |  |
| Voltage | 100ppm/℃ |  |
| Current | 200ppm/℃ |  |
| Display |  |  |
| Display Type | 2.8-inch Color LCD Display (IPS Panel) |  |
| Resolution | 240 × 320 pixels |  |
| Color | 65536 color,TFT |  |
| Environment |  |  |
| Temperature | Working Temperature:0℃～40℃ Storage Temperature:-20℃～60℃ |  |
| Relative Humidity | ≤90%RH; no condensation |  |
| Height | 2,000 meters |  |
| Cooling | Fan cooling, temperature intelligent speed control |  |
| Others |  |  |
| Communication Port | USB, LAN(optional), DB9:RS232(optional), compatible with the SCPI communication protocol |  |
| Dimension | 82mm(W)× 142mm(H)×226mm(D) |  |
| Weight | Approx. 1.6kg |  |

V1.0.0