SPE80 Series Single Channel Power Supply 
Technical Specification
The instrument must be operated continuously for more than 30 minutes at the specified operating temperature to achieve the following specifications:
	Model
	SPE6205
	SPE8105
	SPE8205

	Input Characteristics                                                                                      

	Supply Voltage
	100V-240V

	Input Frequency
	45~65Hz

	Input Fuse
	250V,F10A

	Rated Output                                                                                        

	Voltage
	0~60V
	0~80V
	0~80V

	Current
	0~20A
	0~10A
	0~20A

	Power
	500W
	500W
	500W

	Regulation	(CV)				

	Load
	≤20mV

	Line
	≤20mV

	Regulation	(CC)			

	Load
	≤10mA

	Line
	≤10mA

	Ripple & Noise (Noise bandwidth 20MHz, ripple bandwidth 1MHz, connect 10uF electrolytic capacitor in parallel with 0.1uF ceramic capacitor to the output terminal for testing)

	Voltage (Vp-p)
	≤50mVp-p
	≤50mVp-p
	≤50mVp-p

	Current (rms)
	≤15mA
	≤15mA
	≤15mA

	Setting Resolution

	Voltage
	10mV
	10mV
	10mV

	Current
	1mA
	1mA
	1mA

	Readback Resolution

	Voltage
	10mV
	10mV
	10mV

	Current
	1mA
	1mA
	1mA

	Setting Accuracy

	Voltage
	≤0.05%±20mV

	Current
	≤0.05%±10mA
	≤0.05%±10mA
	≤0.05%±20mA

	Readback Accuracy

	Voltage
	≤0.05%±20mV

	Current
	≤0.05%±10mA
	≤0.05%±10mA
	≤0.05%±20mA

	Response time

	Transient recovery time
(10%~90% rated load)
	≤5.0ms

	Protective function

	OVP
	0~63V
	0~85V
	0~85V

	OCP
	0~21A
	0~11A
	0~21A

	OTP
	85℃
	85℃
	85℃

	Temperature coefficient of output

	Voltage
	100ppm/℃

	Current
	200ppm/℃

	Temperature coefficient of readback value

	Voltage
	100ppm/℃

	Current
	200ppm/℃

	Display

	Display Type
	2.8-inch color LCD display (IPS screen)

	Resolution
	240 × 320 pixels

	Color
	65536 color,TFT

	Environment

	Temperature
	Working Temperature: 0℃～40℃
Storage Temperature: -20℃～60℃

	Relative Humidity
	≤90%RH; no condensation

	Height
	2,000 meters

	Cooling
	Fan cooling, temperature intelligent speed control

	Other

	Communication Port
	USB communication, LAN communication (optional), compatible with SCPI

	Dimension
	82mm（Width）× 142mm（Height）×226mm（Depth）

	Weight
	Approx. 1.5 kg


[image: barcode1]

[bookmark: _GoBack]								                                        	                                                        V1.0.1
image1.png
7007020100197




