SPM Series Source Technical Specification

The instrument must be operated continuously for more than 30 minutes at the specified operating temperature to achieve the following specifications:

Power Supply

| Model | SPM3051 | SPM6053 | SPM3103 | SPM6103 | SPM8104 | SPM8205 |
| --- | --- | --- | --- | --- | --- | --- |
| Input Characteristics |  |  |  |  |  |  |
| Supply Voltage | 110Vac±10% or 220Vac±10%, please supply power according to the mark on the left of the power input socket on the rear panel of the instrument |  |  |  | 100V ~ 240V |  |
| Input Frequency | 50~60Hz | 50~60Hz | 50~60Hz | 50~60Hz | 45~65Hz | 45~65Hz |
| Input Fuse | 250V,F5A | 250V,F10A | 250V,F10A | 250V,F10A | 250V,F10A | 250V,F10A |
| Rated Output |  |  |  |  |  |  |
| Voltage | 0~30V | 0~60V | 0~30V | 0~60V | 0~80V | 0~80V |
| Current | 0~5A | 0~5A | 0~ 10A | 0~ 10A | 0~10A | 0~20A |
| Power | 150W | 300W | 300W | 300W | 400W | 500W |
| Regulation(CV) |  |  |  |  |  |  |
| Load | ≤30mV | ≤30mV | ≤30mV | ≤30mV | ≤20mV | ≤20mV |
| Line | ≤20mV | ≤20mV | ≤20mV | ≤20mV | ≤20mV | ≤20mV |
| Regulation(CC) |  |  |  |  |  |  |
| Load | ≤30mA | ≤30mA | ≤30mA | ≤30mA | ≤10mA | ≤10mA |
| Line | ≤20mA | ≤20mA | ≤20mA | ≤20mA | ≤10mA | ≤10mA |
| Ripple & Noise (Noise bandwidth 20MHz, ripple bandwidth 1MHz, connect 10uF electrolytic capacitor in parallel with 0.1uF ceramic capacitor to the output terminal for testing) |  |  |  |  |  |  |
| Voltage (Vp-p) | ≤30mV | ≤30mV | ≤30mV | ≤50mV | ≤50mVp-p | ≤50mVp-p |
| Voltage (rms) | ≤3mV | ≤3mV | ≤3mV | ≤5mV | ≤15mA | ≤15mA |
| Setting Resolution |  |  |  |  |  |  |
| Voltage | 10mV | 10mV | 10mV | 10mV | 10mV | 10mV |
| Current | 1mA | 1mA | 1mA | 1mA | 1mA | 1mA |
| Readback Resolution |  |  |  |  |  |  |
| Voltage | 10mV | 10mV | 10mV | 10mV | 10mV | 10mV |
| Current | 1mA | 1mA | 1mA | 1mA | 1mA | 1mA |
| Setting Accuracy |  |  |  |  |  |  |
| Voltage | ≤0.1%±20mV | ≤0.1%±20mV | ≤0.1%±20mV | ≤0.1%±30mV | ≤0.05%±20mV | ≤0.05%±20mV |
| Current | ≤0.1%±10mA | ≤0.1%±10mA | ≤0.1%±10mA | ≤0.1%±10mA | ≤0.05%±10mA | ≤0.05%±20mA |
| Readback Accuracy |  |  |  |  |  |  |
| Voltage | ≤0.1%±20mV | ≤0.1%±20mV | ≤0.1%±20mV | ≤0.1%±30mV | ≤0.05%±20mV | ≤0.05%±20mV |
| Current | ≤0.1%±10mA | ≤0.1%±10mA | ≤0.1%±10mA | ≤0.1%±10mA | ≤0.05%±10mA | ≤0.05%±20mA |
| Response time |  |  |  |  |  |  |
| Transient recovery time (10%~90%rated load) | ≤1ms |  |  |  | ≤5.0ms |  |
| Protective function |  |  |  |  |  |  |
| OVP | 0~31V | 0~61V | 0~31V | 0~61V | 0~85V | 0~85V |
| OCP | 0~5.1A | 0~5.1A | 0~ 10.1A | 0~ 10.1A | 0~11A | 0~21A |
| OTP | 85℃ | 85℃ | 85℃ | 85℃ | 85℃ | 85℃ |
| Temperature coefficient of output |  |  |  |  |  |  |
| Voltage | 100ppm/℃ |  |  |  |  |  |
| Current | 200ppm/℃ |  |  |  |  |  |
| Temperature coefficient of readback value |  |  |  |  |  |  |
| Voltage | 100ppm/℃ |  |  |  |  |  |
| Current | 200ppm/℃ |  |  |  |  |  |

Multimeter

| Characteristic | Description |  |
| --- | --- | --- |
| Display | 4½ bit (Max. reading 20000) |  |
| Measurement | Voltage, Current, Resistance, Capacitance, Diode, Continuity |  |
| Input impendance | ≥10 MΩ |  |
| Diode | 0~2V |  |
|  | Range | Accuracy |
| DC Voltage | 200.00mV | ±(0.3%+10dig) |
|  | 2.0000V/20.000V/200.00V | ±(0.3%+5dig) |
|  | 1000.0V |  |
| AC Voltage | 200.00mV /2.0000V/20.000V/200.00V | ±(0.8%+10dig) |
|  | 750V | ±( 1.0%+10dig) |
| DC Current | 200.00mA | ±(0.8%+10dig) |
|  | 10.000A | ±(2.5%+10dig) |
| AC Voltage | 200.00mA | ±(1.0%+10dig) |
|  | 10.000A | ±(2.8%+10dig) |
| Resistance | 200.00Ω/2.0000kΩ/20.000kΩ/ 200.00kΩ/2.0000MΩ | ±(0.8%+10dig) |
|  | 20.000MΩ | ±(1.0%+3dig) |
|  | 100.00MΩ | ±(5.0%+10dig) |
| Capacitance[1](F) | 20.000nF/200.00nF/2.0000uF/ 20.000uF/200.00uF/2.0000mF | ±(3.0%+10dig) |
| Frequency response(Hz) | （40 - 1000）Hz |  |

- The accuracy guarantee range for AC voltage/current and capacitance measurements is from 5% to 100% of the range.

Others

| Display |  |
| --- | --- |
| Display Type | 2.8inches color LCD |
| Resolution | 240 × 320 pixels |
| Color | 65536 color,TFT |
| Environment |  |
| Temperature | Working Temperature:0℃～40℃ Storage Temperature:-20℃～60℃ |
| Relative Humidity | ≤90%RH; no condensation |
| Height | 2,000 meters |
| Cooling | Fan cooling, temperature intelligent speed control |
| Other |  |
| Communication Port | Provides three optional interfaces: USB, LAN(Optional), and RS232(Optional), compatible with the SCPI communication protocol. |
| USB Charging Port | 5V/ 1A charging |
| Dimension | 82mm(W)× 142mm(H)×226mm(D) |
| Weight | Approx. 1.5 kg |

- V1.1.1